## High Flow Oxygen - NICU
High Flow Oxygen Delivery Unit (NICU) |  |  |  |  | 37 | <-- Total score
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 39557 Oxygen Delivery Systems, High Flow Total Quantity: 39 | Equipment Location (Region): Quantity SIHREN a. Indonesia (Nationwide): 39 units | Description: High-flow oxygen therapy systems are commonly used to treat acute respiratory distress (e.g., bronchiolitis) in spontaneously-breathing neonates and infants and in children and adults (e.g., for chronic obstructive pulmonary disease [COPD]). Units designed to deliver a high-flow (typically up to 8 L/min for neonatal patients or above 15 L/min for adults) of heated and humidified air or mixture of air and oxygen to a patient. These units usually consist of a flow generator, active heated humidifier, breathing circuit, and nasal cannula. High-flow oxygen delivery units have adjustable temperature and flow settings and can provide air/oxygen flows up to 60L/min. Some units also have tracheostomy or mask interfaces for delivery of the air/oxygen flow.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder to provide product brochure in English version. | Enter Responses Here!
Nombor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Delivery of High Flow Oxygen | Required |  |  | Enter Responses Here!
Patient Type - Premature baby to Neonate | Required for neonate less than 1 kg |  |  | Enter Responses Here!
Patient Type - Newborn Infant | Preferred |  |  | Enter Responses Here! | 2
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Flow Range, LPM | Required - Must have Flow rate range of minimum 1 to 8 lpm for neonates |  |  | Enter Responses Here!
Flow Range, LPM | Preferred - Flow rate range 1-25 lpm for Newborn Infant |  |  |  | 2
Flow Range, LPM | Preferred - Flow rate range from 1 lpm to 40 lpm preferred to accomodate a wider range of patients (e.g. neonates) |  |  | Enter Responses Here! | 2
Temperature Settings | Required - At least 2 settings available within the range of 34 to 37 degrees Celcius |  |  | Enter Responses Here!
Temperature adjustment Range | Preferred - Fully adjustable within the range of 31 - 40 degrees Celcius |  |  | Enter Responses Here! | 2
Water level detection | Required - Visible low water level line on the chamber |  |  | Enter Responses Here! | 2
FiO2 Range % | 21 to 100% |  |  | Enter Responses Here!
Alarms
Temperature Alarm | Required |  |  | Enter Responses Here!
Low Water Level Alarm | Preferred |  |  | Enter Responses Here! | 1
Oxygen Alarm | Required |  |  | Enter Responses Here!
Low Battery | Required if internal battery is available |  |  | Enter Responses Here!
Blocked Tube | Preferred |  |  | Enter Responses Here! | 1
Power Out | Preferred |  |  | Enter Responses Here! | 1
Breathing Tube fault | Preferred |  |  | Enter Responses Here! | 1
Flow Error | Preferred |  |  | Enter Responses Here! | 1
Gas Supply | Preferred |  |  | Enter Responses Here! | 1
Configuration
Display | Required |  |  | Enter Responses Here!
Display Type | Bidder to Specify; LED ; LCD ; OLED ; Touchscreen |  |  | Enter Responses Here!
Mounting Type | Bidder to Specify; Pole stand; Roll stand; Trolley ; Tabletop |  |  | Enter Responses Here!
Type of Disinfectant used for Infection Control | Free of corrosive ingredients |  |  | Enter Responses Here!
Compatible Agents | Bidder to specify |  |  | Enter Responses Here!
Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Heated Humidifier | Required |  |  | Enter Responses Here!
Flow Generator | Required |  |  | Enter Responses Here!
Air/ Oxygen Blender | Required |  |  | Enter Responses Here!
O2 inlet kit | Preferred |  |  | Enter Responses Here! | 2
START-UP CONSUMABLES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Startup High Flow Breathing Circuit | Bidder to supply quantity for 100 typical patient procedures |  |  | Enter Responses Here!
Startup Nasal Cannula | Bidder to supply quantity for 100 typical patient procedures, to include various sizes |  |  | Enter Responses Here!
High Flow Breathing Circuit Needed | Bidder to supply quantity for 100 typical patient procedures, to include various sizes |  |  | Enter Responses Here!
Nasal Cannula Needed | Bidder to supply quantity for 100 typical patient procedures, to include various sizes |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Weight of device (in kg) | Bidder to specify |  |  | Enter Responses Here!
MECHANICAL & ELECTRICAL (M&E) REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, KWH | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Battery Powered | Required |  |  | Enter Responses Here!
Battery Type | Preferred; Rechargeable |  |  | Enter Responses Here! | 2
Battery Operating Time, hr | > 30 minutes; Bidder to specify |  |  | Enter Responses Here!
Battery Recharging Time, hr | Bidder to specify |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 2
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | N/A |  |  | Enter Responses Here!
Compatibility with DICOM (Digital Imaging and Communications in Medicine) | N/A |  |  | Enter Responses Here!
Type of Networking/Connectivity Interface | N/A |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | N/A |  |  | Enter Responses Here!
Data Port | Data port must be activated (e.g: RS232, USB) to allow download of data logs from device and service |  |  | Enter Responses Here!
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware updates and upgrades at no cost |  |  | Enter Responses Here!
Remote Support Capability | N/A |  |  | Enter Responses Here!
Cybersecurity Considerations
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | N/A |  |  | Enter Responses Here!
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here! | 1
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
MEDICAL GASES REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Temperature Requirements | The areas temperature ranges from 22-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed. |  |  | Enter Responses Here!
Humidity Requirements | The area's relative humidity ranges between 70% and 90%.Equipment must be operational in that range. If it is not please specify the range needed. |  |  | Enter Responses Here!
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 14
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of High Flow Oxygen Delivery Unit (NICU)
Include description and pricing for all accessories (e.g., transducers, monitors, carts, etc.) required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! |  | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of High Flow Oxygen Delivery Unit (NICU)
Include description and pricing for all consumable items (e.g., Ultrasound gel, etc.) required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Vapotherm | Consumables | Low Flow Disp Patient Circuit (5Pc/Box) |  | 1 BX = 5 PC | 610 | 610
Vapotherm | Consumables | High Flow Disp Patient Circuit (5Pc/Box) |  | 1 BX = 5 PC | 610 | 610
Vapotherm | Consumables | Premature Cannula (25Pc/Box) |  | 1 BX = 25 PC | 142 | 142
Vapotherm | Consumables | Neonatal Cannula (25Pc/Box) |  | 1 BX = 25 PC | 142 | 142
Vapotherm | Consumables | Infant Cannula (25Pc/Box) |  | 1 BX = 25 PC | 142 | 142
Vapotherm | Consumables | Intermediate Infant Cannula (25Pc/Box) |  | 1 BX = 25 PC | 142 | 142
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 1788 | 1788
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 1788

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of High Flow Oxygen Delivery Unit (NICU)
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
Vapotherm | Precission Flow | 1 year | 3x | Labour Only (include 1x filter repleacement) | 31487300.625 | 34636030.6875 | 38099633.7563 | 41909597.1319 | 46100556.8451
Vapotherm | Precission Flow | 1 year | 3x | All Part (all part warranty, include 1x O2 Cell, 1x Filter Replacement, 1x Battery) | 59318011.5 | 65249812.65 | 71774793.915 | 78952273.3065 | 86847500.6372