## Neonate Inv. Patient Monitor
Neonate Invasive Patient Monitor (NICU) |  |  |  |  | 55
Bidders should complete this form to provide responses about their device which meet or exceed these required specifications, and shall provide supporting documents, where applicable.
UMDNS: 15791 Monitors, Physiologic, Neonatal/Infant, Bedside Total Quantity: 40 | Equipment Location (Region): Quantity SIHREN a. Indonesia (Nationwide): 40 units | Description: Physiologic monitors designed for continuous checking of vital physiologic parameters of newborns and infants. These monitors should include sensors with appropriate size and design for infant use; they may also include capabilities to assess hemoglobin oxygen saturation (SpO2) through transcutaneous sensors that measure both transcutaneous oxygen (tcPO2) and transcutaneous carbon dioxide (tcPCO2) saturation. Neonatal monitors are used for newborn and premature infants, especially those that are under critical care.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
For monitoring of neonate physiologic data | For neonate, especially those that are under critical care. |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Patient Type | Neonate |  |  | Enter Responses Here!
Configuration | Bedside portable |  |  | Enter Responses Here!
To be mounted on Gypsum wall | Bidder to provide suitable wall-mount bracket or support |  |  | Enter Responses Here!
Type | Configured with built-in modules and additional modules such as: |  |  | Enter Responses Here!
Required Parameters | a) Electrocardiography (ECG) |  |  | Enter Responses Here!
| b) Non-invasive blood pressure (NIBP) |  |  | Enter Responses Here!
| c) Invasive blood pressure (IBP) |  |  | Enter Responses Here!
| c) Pulse oximetry (SpO2) |  |  | Enter Responses Here!
| d) Heart rate (HR) |  |  | Enter Responses Here!
| e) Temperature |  |  | Enter Responses Here!
| f) Respiration |  |  | Enter Responses Here!
Module Rack | Transportable and compatible with proposed physiologic monitors in Operating Rooms, if modular configuration available |  |  | Enter Responses Here!
a) Electrocardiography (ECG) | 12-leads interpretation |  |  | Enter Responses Here!
Arrhythmia Detection | Required for full arrhythmia |  |  | Enter Responses Here!
No. of Leads Analyzed | Bidder to specify |  |  | Enter Responses Here!
ST Analysis | Required |  |  | Enter Responses Here!
No. of Leads Analyzed | Bidder to specify |  |  | Enter Responses Here!
ST Segment Trending | Required |  |  | Enter Responses Here!
Length of Time, hr | ≥24 hours |  |  | Enter Responses Here!
Length of Time, hr | ≥ 48 hours |  |  | Enter Responses Here! | 3
Full Disclosure (FD) | Ability to continuously store beat-to-beat physiologic waveforms over an extended period of time |  |  | Enter Responses Here! | 3
Length of Time, hr | ≥24 hours |  |  | Enter Responses Here!
Length of Time, hr | ≥ 48 hours |  |  | Enter Responses Here! | 3
Leads Disconnected Alarm | Required |  |  | Enter Responses Here!
b) Non-invasive blood pressure (NIBP) | For neonate |  |  | Enter Responses Here!
Measuring Method | Oscillometric or auscultatory Bidder to specify |  |  | Enter Responses Here!
Systolic Measuring Range, mm Hg | 20-120 mmHg for neonates |  |  | Enter Responses Here!
Diastolic Measuring Range, mm Hg | 10-100 mmHg for neonates |  |  | Enter Responses Here!
Mean Arterial Pressure (MAP) Measuring Range, mm Hg | 30-110 mmHg for neonates |  |  | Enter Responses Here!
Cuff Inflation Rate, mmHg/sec | Should not be greater than 80 mmHg/sec |  |  | Enter Responses Here!
Cuff Deflation Rate, mmHg/sec | Should be 2 to 3 mmHg/sec, depending on pulse rate |  |  | Enter Responses Here!
Measurement Intervals, minutes | User selectable; ≥5 choices |  |  | Enter Responses Here!
| Continuous Measurement |  |  | Enter Responses Here!
c) Invasive blood pressure (IBP) | For neonate |  |  | Enter Responses Here!
No. of channels | 3 is preferred. Bidder to specify number of channels |  |  | Enter Responses Here! | 3
c) Pulse oximetry (SpO2) | For neonate |  |  | Enter Responses Here!
Probe Type | Reusable, Foot (preferred) or forehead probe. Bidder to specify |  |  | Enter Responses Here!
Technology Type | Bidder to specify |  |  | Enter Responses Here!
d) Heart Rate (HR) | Required |  |  | Enter Responses Here!
Pulse Measuring Range, bpm | 30-180 bpm for neonates |  |  | Enter Responses Here!
e) Temperature | For neonate |  |  | Enter Responses Here!
Probe Type | Reusable, Skin probe |  |  | Enter Responses Here!
f) Respiration | Required |  |  | Enter Responses Here!
Measuring Method | Impedance pneumography to identify apnea |  |  | Enter Responses Here!
Waveform Displayed | Required |  |  | Enter Responses Here!
Threshold Control | Automatic or Manual setting for the minimum impedance necessary to count respiration |  |  | Enter Responses Here!
Display Type | LCD is preferred |  |  | Enter Responses Here!
Size (diagonal), in | ≥12 inches |  |  | Enter Responses Here!
Traces | 6 traces to be viewed simultaneously on the screen |  |  | Enter Responses Here!
Traces | ≥ 8 traces to be viewed simultaneously on the screen |  |  | Enter Responses Here! | 3
Type of User Interface | Touchcreen, Quick access key, Trim knob preferred |  |  | Enter Responses Here! | 2
Recorder for Documentation | Required |  |  | Enter Responses Here!
Type of Alarm | Audible and Visual |  |  | Enter Responses Here!
Silence Active Alarm | User configurable |  |  | Enter Responses Here!
Silence Duration, sec | Bidder to specify |  |  | Enter Responses Here!
Alarm Suspension | Preferred; If the monitor allows for alarm suspension, there should be a clear indication on the display and remains visible as long as alarms are suspended |  |  | Enter Responses Here! | 2
Suspend Duration, sec | Bidder to specify |  |  | Enter Responses Here!
Alarm Latching | Alarm continues even if the alarm condition self-corrected until the alarm is manually acknowledged |  |  | Enter Responses Here!
Alarm History Tracking | Required |  |  | Enter Responses Here!
Type of Device Alarm Conditions | a) Cuff leak or Cuff disconnection |  |  | Enter Responses Here!
| b) Hose leak |  |  | Enter Responses Here!
| c) Inflation/Deflation error |  |  | Enter Responses Here!
| d) Failure to take successful reading |  |  | Enter Responses Here!
| e) Low battery |  |  | Enter Responses Here!
| f) Arrhythmia detected |  |  | Enter Responses Here!
| g) Measured neonatal monitoring parameters reaches the limit set |  |  | Enter Responses Here!
| h) Apnea detection |  |  | Enter Responses Here!
Alarm Setting | Bidder to specify |  |  | Enter Responses Here!
Low Systolic BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
High Systolic BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
Low Diastolic BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
High Diastolic BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
Low Mean BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
Low Respiratory Rate, bpm | Bidder to specify the setting range, bpm |  |  | Enter Responses Here!
High Respiratory Rate, bpm | Bidder to specify the setting range, bpm |  |  | Enter Responses Here!
Low IBP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
High IBP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
Low Mean BP, mm Hg | Bidder to specify the setting range, mmHg |  |  | Enter Responses Here!
Low Heart Rate, bpm | Bidder to specify the setting range, bpm |  |  | Enter Responses Here!
High Heart Rate, bpm | Bidder to specify the setting range, bpm |  |  | Enter Responses Here!
Low Temperature, T | Bidder to specify the setting range, ⁰C |  |  | Enter Responses Here!
Low Temperature, T | Bidder to specify the setting range, ⁰C |  |  | Enter Responses Here!
Low Pulse Oximetry, % | Bidder to specify the setting range, % |  |  | Enter Responses Here!
Alarm Priority/Level | Critical, High, Medium, Low |  |  | Enter Responses Here!
User Defined/Configured | Required |  |  | Enter Responses Here!
Colour Coded | Required |  |  | Enter Responses Here!
Volume/Tones | Escalating and varies according to priority of alarm. Volume may be muted as well. |  |  | Enter Responses Here!
Type of Disinfectant used for Infection Control | Free of corrosive ingredients |  |  | Enter Responses Here!
Compatible Agents | Bidder to specify |  |  | Enter Responses Here!
Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) ECG trunk cable and leadwire (3, 5 or 6 leads) | 2 pc per unit |  |  | Enter Responses Here!
b) Dual NIBP hose | 2 pc per unit |  |  | Enter Responses Here!
Enter Responses Here!
c) NIBP cuff, size neonate | 2 pc per unit |  |  | Enter Responses Here!
d) NIBP cuff, size thigh neonate | 2 pc per unit |  |  | Enter Responses Here!
e) SpO2 connection cable | 2 pc per unit |  |  | Enter Responses Here!
f) SpO2 Probe, size neonate | 2 pc per unit |  |  | Enter Responses Here!
g) Skin temperature probe | 2 pc per unit |  |  | Enter Responses Here!
i) Patient Data Recording module | 1 pc per unit |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
START-UP CONSUMABLES, if any | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) ECG electrode, size neonate | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
b) Arterial Catheter-Transducer Set (IBP) | Bidder to supply quantity for 100 typical patient procedures. To include various neonate sizes. |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dimension, H x W x D, cm | Bidder to specify | - |  | Enter Responses Here!
Weight, kg | Bidder to specify | - |  | Enter Responses Here!
MECHANICAL & ELECTRICAL (M&E) REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, W | Bidder to specify | - |  | Enter Responses Here!
Battery Operation | Required |  |  | Enter Responses Here!
Battery Type | Rechargeable |  |  | Enter Responses Here!
Battery operating time, min | Bidder to specify |  |  | Enter Responses Here!
Battery recharging time, hr | Bidder to specify |  |  | Enter Responses Here!
Battery Indicator | Required |  |  | Enter Responses Here!
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | Required; e.g: HIS, RIS, EMR, etc |  |  | Enter Responses Here!
Compatibility with DICOM (Digital Imaging and Communications in Medicine) | N/A |  |  | Enter Responses Here!
Networked to Central Station Monitor | Bidder to install connectivity |  |  | Enter Responses Here!
Bedside-to-bedside Communication | Required |  |  | Enter Responses Here!
Type of Networking/Connectivity Interface | Wireless or LAN-Port, or both Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
Data Port | Data port must be activated (e.g: RS232, USB) to allow download of data logs from device and service |  |  | Enter Responses Here!
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
Type of Data Transferred | All data monitored, alarms history, etc. Bidder to specify |  |  | Enter Responses Here!
Method of Data Transferred | Real-time and/or batch Bidder to specify |  |  | Enter Responses Here!
Peripheral Device Interface | Required |  |  | Enter Responses Here!
Ventilators | Bidder to specify compatible brand/model |  |  | Enter Responses Here!
Infusion Pumps | Bidder to specify compatible brand/model |  |  | Enter Responses Here!
Syringe Pumps | Bidder to specify compatible brand/model |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware upgradeable at no cost |  |  | Enter Responses Here!
Remote Support Capability | Preferred; Off site diagnosis of technical issues and remote fixes. |  |  | Enter Responses Here! | 3
Cybersecurity Considerations
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | Preferred |  |  | Enter Responses Here! | 2
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here! | 1
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Does the equipment include any AI Capability? | Bidder to specify AI Capabilities. |  |  | Enter Responses Here! | 3
What is the AI capability intended to assist with? (e.g., workflow automation, diagnosis assistance, patient scheduling, other) | Bidder to Specify |  |  | Enter Responses Here!
Does the AI Capability have High or Low impact on Patient Care? | Bidder to Specify |  |  | Enter Responses Here!
Is this equipment that incorporates AI capabilities generally classified as high-risk or low-risk? | Bidder to specify |  |  | Enter Responses Here!
Are the AI capabilities embedded in the device or require of third party applications? | Bidder to specify |  |  | Enter Responses Here!
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Does the equipment have telemedicine capability or can it be used in telemedicine applications? | Bidder to specify telemedicine capabilities |  |  | Enter Responses Here! | 3
Asynchronous telemedicine capabilities. Store-and-forward technology is asynchronous in that it allows for the images or data to be interpreted at a later time and eliminates the need for the patient and physician to be in the same room or venue. (i.e., through store and forward, in which patient information, such as an image or data, "stored" in the originating facility's (i.e., where the patient is) computer system is sent securely ("forwarded") to the telehealth provider at the distant site (i.e., where the provider is) for review.) | Bidder to specify |  |  | Enter Responses Here!
Remote patient monitoring capabilities. Remote patient monitoring, in which vital sign information (e.g., blood pressure) and other data (e.g., blood glucose levels) are collected by monitoring devices, usually at home, and transferred to the patient's physician. This technology is used often in the care of patients with chronic conditions. | Bidder to specify |  |  | Enter Responses Here!
Mobile health (mHealth) integration capabilities. Mobile health (mHealth), in which healthcare and public health information are communicated through mobile devices. The information may include general educational information, targeted texts, and notifications about disease outbreaks. | Bidder to specify |  |  | Enter Responses Here!
Other telemdicine capabilities. | Bidder to specify |  |  | Enter Responses Here!
Is telemedicine offered through a 3rd party vendor? | Bidder to specify |  |  | Enter Responses Here!
Video Capabilities (if available)
Are video capabilities needed? If so, specify if they are included in the solution or offered through 3rd party add-on? | Bidder to specify |  |  | Enter Responses Here! | 2
If video capabilities are available, specify the resolution | Bidder to specify |  |  | Enter Responses Here!
Is night vision available? | Bidder to specify |  |  | Enter Responses Here!
Does the system offer pan/tilt features? | Bidder to specify |  |  | Enter Responses Here!
Is video zoom available? | Bidder to specify |  |  | Enter Responses Here!
Audio capabilities (if available)
Is audio functionality available? | Bidder to specify |  |  | Enter Responses Here! | 1
Specify audio features | Bidder to specify |  |  | Enter Responses Here!
Is echo cancellation available? | Bidder to specify |  |  | Enter Responses Here!
Is noise cancellation available? | Bidder to specify |  |  | Enter Responses Here!
MEDICAL GASES REQUIREMENTS 医用气体要求 | REQUIREMENTS 要求 | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
MRI compatible medical gas panel and waveguides within RF Cabin | Bidder to provide installation |  |  | Enter Responses Here!
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Temperature Requirements | The areas temperature ranges from 22-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed. |  |  | Enter Responses Here!
Humidity Requirements | The area's relative humidity ranges between 70% and 90%.Equipment must be operational in that range. If it is not please specify the range needed. |  |  | Enter Responses Here!
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 21
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Neonate Invasive Patient Monitor (NICU)
Include description and pricing for all accessories (e.g., <insert examples>, etc.) required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
GE Healthcare | ECG | ECG TRUNK CABLE, NEONATAL DIN 3-LEAD, AHA, 3.6 M/12 FT | Enter Response Here! | 1 | 8000000 | 8000000
GE Healthcare | ECG | Neonatal cloth radiolucent electrodes with pre-attached DIN |  | 1 | 30000000 | 30000000
GE Healthcare | Cuff | NEO-SNAP Airhose, 12 ft (3.6M) Care Hose Rectangle Connector |  | 1 | 7000000 | 7000000
GE Healthcare | Cuff | SOFT-CUF Assortment Kit, Neonate 2T NEO-SNAP 20/PK |  | 1 | 10000000 | 10000000
InnoQ | InnoQ | INNOQ SPO2 SENSOR ADAPTOR CABLE - G (2.2M) |  | 1 | 1144200 | 1144200 | https://e-katalog.lkpp.go.id/katalog/produk/detail/61821854?lang=id&type=general
GE Healthcare | Temp | TEMPERATURE CABLE, 400 SERIES PROBES, DUAL, 0.5 M/1.7 FT. |  | 1 | 4950000 | 4950000
GE Healthcare | Temp | REUSABLE TEMPERATURE PROBE - SKIN |  | 1 | 7875000 | 7875000
GE Healthcare | ECG | Trunk - 3/5-Lead AHA - 3.6M |  | 1 | 6000000 | 6000000
GE Healthcare | ECG | Trunk - 12-Lead AHA - 3.6M |  | 1 | 12000000 | 12000000
GE Healthcare | ECG | Leadwire - 5-Lead AHA - 74CM |  | 1 | 7000000 | 7000000
GE Healthcare | ECG | Leadwire - V2-6 AHA - 74CM |  | 1 | 6500000 | 6500000
InnoQ | InnoQ | INNOQ SPO2 SENSOR ADAPTOR CABLE - G (2.2M) |  | 1 | 300000 | 300000 | https://e-katalog.lkpp.go.id/katalog/produk/detail/61808183?lang=id&type=general
GE Healthcare | Cuff | NIBP Hose - Adult Dinaclick, 3.6M |  | 1 | 2000000 | 2000000
InnoQ | Cuff | INNOQ REUSABLE NIBP CUFF (BLADDERLESS), 2-TUBES, D-G, CHILD |  | 1 | 284000 | 284000 | https://e-katalog.lkpp.go.id/katalog/produk/detail/62246348?lang=id&type=general
0
0
|  | TOTAL |  | 14 | 103053200 | 103053200
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 103053200

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of Neonate Invasive Patient Monitor (NICU)
Include description and pricing for all consumable items (e.g.,<insert examples> etc.) required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Vygon | Arterial Line | Arterial Leaderflex | 42000000-AKN-000684343 | 1 | 795000 | 795000
Bioptimal | Monitoring Kit | Monitoring Kit Double lumen | 4815000210-FKS-005691193 | 1 | 358000 | 358000
Bioptimal | CVC | BIOPTIMAL CVC DOUBLE LUMEN: 4FR X 16CM WITH KINK RESISTANCE GUIDEWIRE, STRAIGHT NEEDLE (CVN-402-16) |  | 1 | 515000000 | 515000000
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 3 | 516153000 | 516153000
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 516153000

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Neonate Invasive Patient Monitor (NICU)
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
GE Healthcare | GE B125M PATIENT MONITOR 6P +1512L ECG |  | 3 | Preventive maintenance (visual device and acc.check, function test and cleaning) and Calibration. | 12149580 | 12149580 | 12149580 | 12149580 | 12149580