## Ventilator (Neonate - NICU)
Ventilator (Neonate - NICU) |  |  |  |  | 78 | <-- Total score
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS 14361 Ventilators, Intensive Care, Neonatal/Pediatric Total Quantity: 414 | Equipment Location (Region): Quantity SIHREN a. West Indonesia: 155 units b. Central Indonesia: 75 units c. East Indonesia: 184 units | Description: Intensive care ventilators designed for neonatal and/or pediatric respiratory support; most are time-cycled pressure-control devices. These ventilators typically include flexible, small-bore (e.g., 1 cm diameter) tubing and continuous positive airway pressure (CPAP) and/or positive end-expiratory pressure (PEEP) controls. CPAP is useful for infants with restrictive lung disease or decreased lung compliance and alveolar collapse (infants with hyaline membrane disease); PEEP maintains lung volume and prevents alveolar collapse. Neonatal/pediatric intensive care ventilators are used for preterm and critically ill infants with respiratory failure.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Provide temporary or permanent ventilatory support for patients who cannot breathe on their own or who require assistance maintaining adequate ventilation. | Dedicated ventilator for patients with weight 3 to 20kg |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Type | Mobile, with swivel castors and locks |  |  | Enter Responses Here!
Controls/Setting Ranges
Tidal Volume, mL | 2 - 300mL |  |  | Enter Responses Here!
Respiratory Rate, breaths/min | 1 - 150 breaths/min |  |  | Enter Responses Here!
Trigger Mechanism | Either pressure or flow triggering. Bidder to specify |  |  | Enter Responses Here!
Fraction of Inspired Oxygen, FiO2, % | Adjustable from 21% to 100% |  |  | Enter Responses Here!
Inspiratory Flow Rate, L/min | Adjustable, 0 - 30 L/min approx. |  |  | Enter Responses Here!
Flow Pattern/Waveform Adjustment | Required |  |  | Enter Responses Here!
Inspiratory Pressure, cmH2O | 0 - 80 cmH2O approx. |  |  | Enter Responses Here!
Ratio of Inspiratory Time to Expiratory Time (IE Ratio) | Bidder to specify |  |  | Enter Responses Here!
Sigh Breath Function | Preferred |  |  | Enter Responses Here! | 3
PEEP/CPAP, cmH2O | 0 - 25 cmH2O |  |  | Enter Responses Here!
PEEP/CPAP, cmH2O | 0 - 50 cmH2O |  |  | Enter Responses Here! | 2
Pressure Support, cm H2O | 0 - 60 cmH2O (above PEEP) Capable of providing pressure support during all spontaneous breathing modes. |  |  | Enter Responses Here!
Leak Compensation | Preferred, automatically adjusts delivered flow to compensate for variable leaks. |  |  | Enter Responses Here! | 3
Auto 100% / Increase O2 button | Required |  |  | Enter Responses Here!
O2 Emergency Flow Rate, L/min | 0-10 L/min |  |  | Enter Responses Here!
Control Panel Lock | Required |  |  | Enter Responses Here!
Operating Modes | Different modes to provide either full or partial ventilatory support, depending on patient's condition. |  |  | Enter Responses Here!
a) Pressure Control Ventilation (PCV) Mode | Required; Bidder to specify full list of PC modes, or equivalents |  |  | Enter Responses Here!
b) Volume Control Ventilation (VCV) Mode | Required; Bidder to specify full list of VC modes, or equivalents |  |  | Enter Responses Here!
c) Synchronized Intermittent Mandatory Ventilation Based on Pressure Controlled Ventilation with Pressure Support (SIMV (PC) +PS) Mode | Required; SIMV (PC) +PS or equivalents |  |  | Enter Responses Here!
d) Volume Synchronized Intermittent Mandatory Ventilation (VSIMV) | Required; Bidder to specify full list of VSIMV modes, or equivalents |  |  | Enter Responses Here!
e) Pressure Synchronized Intermittent Mandatory Ventilation (PSIMV) | Required; Bidder to specify full list of PSIMV modes, or equivalents |  |  | Enter Responses Here!
f) Synchronized Intermittent Mandatory Ventilation Based on Volume Controlled Ventilation with Pressure Support (SIMV (VC) + PS) Mode | Required; SIMV (VC) + PS or equivalents |  |  | Enter Responses Here!
g) Synchronized Intermittent Mandatory Ventilation Based on Pressure Regulated Volume Controlled Ventilation with Pressure Support (SIMV (PRVC) + PS) Mode | Required; SIMV (PRVC) + PS or equivalents |  |  | Enter Responses Here!
h) Pressure Regulated Volume Control (PRVC) | Required; Controlled mode of ventilation which combines pressure and volume controlled ventilation. Bidder to specify relevant PRVC modes |  |  | Enter Responses Here!
i) Continuous Positive Airway Pressure (CPAP)/ Pressure Support Ventilation (PSV) | Required; Bidder to specify full list of CPAP/PSV or PS/CPAP, or equivalents |  |  | Enter Responses Here!
j) Non-Invasive Ventilation (NIV/NPPV) | Required. To facilitate short-term use of the ventilator on patients without an artificial airway (i.e., endotracheal tube or tracheostomy tube) |  |  | Enter Responses Here!
NIV Pressure Control | Required |  |  | Enter Responses Here!
NIV Pressure Support | Required |  |  | Enter Responses Here!
k) Bilevel Positive Airway Pressure (BPAP) | Required; Bidder to specify full list of BPAP modes, or equivalents |  |  | Enter Responses Here!
l) Airway Pressure Release Ventilation (ARPV) | Required; Bidder to specify full list of ARPV modes, or equivalents |  |  | Enter Responses Here!
m) Advanced Ventilator Modes | Preferred; Intelligent or optimal breath targeting scheme |  |  | Enter Responses Here! | 3
n) Others Operating Mode | Bidder to specify |  |  | Enter Responses Here! | 3
o) Backup Ventilation for Apnea Capability | Required |  |  | Enter Responses Here!
Patient Monitoring Features
Display Type | LCD touch screen |  |  | Enter Responses Here!
Screen Size, inches | Minimum 10", tiltable to enhance visability |  |  | Enter Responses Here!
Waveform and Scalars | Graphs of pressure, volume and flow vs time |  |  | Enter Responses Here!
Airway Pressure (Paw) | Required |  |  | Enter Responses Here!
Flow | Able to display flow rate |  |  | Enter Responses Here!
Volume | Required |  |  | Enter Responses Here!
Monitored/Displayed Parameters | Real-time display |  |  | Enter Responses Here!
Peak Inspiratory Pressure (PIP) | Required |  |  | Enter Responses Here!
Mean Airway Pressure (MAP) | Required |  |  | Enter Responses Here!
PEEP Pressure | Required |  |  | Enter Responses Here!
Tidal Volume | Required |  |  | Enter Responses Here!
Minute Volume | Required |  |  | Enter Responses Here!
Spontaneous Minute Volume | Required |  |  | Enter Responses Here!
FiO2 (Analyzed %) | Required |  |  | Enter Responses Here!
Respiratory Rate | Required |  |  | Enter Responses Here!
Inspiratory Time | Required |  |  | Enter Responses Here!
Expiratory Time | Required |  |  | Enter Responses Here!
IE Ratio | Required |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
Alarm | Audible and visual |  |  | Enter Responses Here!
Alarm Priority/Level | Critical, High, Medium, Low |  |  | Enter Responses Here!
User Defined/Configured | Required |  |  | Enter Responses Here!
Colour Coded | Required |  |  | Enter Responses Here!
Volume/Tones | Escalating and varies according to priority of alarm |  |  | Enter Responses Here!
Alarm Latching | Alarm continues even if the alarm condition self-corrected until the alarm is manually acknowledged |  |  | Enter Responses Here!
Alarm Silencing | a) Unable to silent unattended high priority alarms |  |  | Enter Responses Here!
| b) If an alarm is silenced, it must reactivate automatically within two minutes if the condition is not corrected. A visual display should remain on to clearly indicate which alarm is disabled |  |  | Enter Responses Here!
Silence Duration, sec | Maximum 120 seconds |  |  | Enter Responses Here! | 1
Patient Alarms | Audible and visual |  |  | Enter Responses Here!
Low/high FiO2 | Required |  |  | Enter Responses Here!
Low minute volume | Required |  |  | Enter Responses Here!
High minute volume | Required |  |  | Enter Responses Here!
Low inspiratory pressure | Required |  |  | Enter Responses Here!
High inspiratory pressure | Required |  |  | Enter Responses Here!
High respiratory rate | Required |  |  | Enter Responses Here!
Continuous high pressure/occlusion | Required |  |  | Enter Responses Here!
High PEEP | Required |  |  | Enter Responses Here!
Loss of PEEP | Required |  |  | Enter Responses Here!
Apnea | Required |  |  | Enter Responses Here!
Inverse IE ratio | Required |  |  | Enter Responses Here!
Breathing circuit disconnect | Required |  |  | Enter Responses Here!
Others | Bidder to specify, if any |  |  | Enter Responses Here!
Equipment Alarms | Audible and visual |  |  | Enter Responses Here!
Gas supply failure | Required |  |  | Enter Responses Here!
Power failure | Required |  |  | Enter Responses Here!
Vent inoperative | Required |  |  | Enter Responses Here!
Low battery | Required |  |  | Enter Responses Here!
Self-diagnostics failure | Required |  |  | Enter Responses Here!
Others | Bidder to specify, if any |  |  | Enter Responses Here!
Report System | Required |  |  | Enter Responses Here!
Data Type | User configurable; All data measured, e.g: ventilator alarms and patient status |  |  | Enter Responses Here!
View Reports on Display | Bidder to specify |  |  | Enter Responses Here!
Sends Reports to Printer | Bidder to specify |  |  | Enter Responses Here!
Save Data to External Media | Bidder to specify |  |  | Enter Responses Here!
Send Data via Network | Bidder to specify |  |  | Enter Responses Here!
System Security | Password protected |  |  | Enter Responses Here!
Nebulizer | Required |  |  | Enter Responses Here!
Humidifier | Required |  |  | Enter Responses Here!
Oxygen-Air Proportional (Blender) | Required |  |  | Enter Responses Here!
Compliant with ISO 80601-2-12 - Medical electrical equipment — Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Built-in/Integrated Turbine Compressor | Preferred; to allow operation without being connected to a oxygen source for FIO2 at 21% |  |  | Enter Responses Here! | 5
Oxygen enrichment feature | Allow high concentration of oxygen to be reached within a set amount of time (1 or 2 mins) |  |  | Enter Responses Here! | 3
Heliox Capability | Preferred; the capability to provide ventilation using helium/oxygen (heliox) mixtures, for patients with obstructive lung disease |  |  | Enter Responses Here! | 2
Lung Diagnostic and Assessment Tools (Loops) | Required; Functions to graph PV loops and FV loops |  |  | Enter Responses Here!
Lung Recruitment Tools and Maneuvers | Required; allows necessary tools for adjusting ventilation parameters in order to perform recruitment maneuvers. This includes adjustable PEEP and inspiratory hold. |  |  | Enter Responses Here!
Lung Mechanics Information and Lung Visualization Tools | Preferred; Provide computed lung-mechanics parameters and graphical representations of lung mechanics should show changes in lung resistance and compliance. |  |  | Enter Responses Here! | 3
Event Log | Data log that includes ventilation setting, setting changes and alarms, and can store up to 72 hours |  |  | Enter Responses Here!
Type of Disinfectant used for Infection Control | Free of corrosive ingredients |  |  | Enter Responses Here!
Compatible Agents | Bidder to specify |  |  | Enter Responses Here!
Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) Air high pressure tubing | 2 pcs per unit |  |  | Enter Responses Here!
b) O2 high pressure tubing | 2 pcs per unit |  |  | Enter Responses Here!
c) Breathing Circuit, Size Neonate | 3 pcs per unit |  |  | Enter Responses Here!
d) Test Lung, Size Neonate | 1 pc per unit |  |  | Enter Responses Here!
e) Reusable Expiratory Air Filter | 3 pcs per unit |  |  | Enter Responses Here!
f) Nebulizer | 1 pc per unit |  |  | Enter Responses Here!
g) Humidifier | 1 pc per unit |  |  | Enter Responses Here!
h) Oxygen-Air Proportioner (Blender) | 1 pc per unit |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
START-UP CONSUMABLES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) Flow Sensor | Bidder to specify quantity for 100 typical patient procedures |  |  | Enter Responses Here!
b) O2 Sensor | Bidder to specify quantity for 100 typical patient procedures |  |  | Enter Responses Here!
c) Face Mask, Size Neonate | Bidder to specify quantity for 100 typical patient procedures |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Weight, kg | Bidder to specify |  |  | Enter Responses Here!
ELECTRICAL REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Main Battery | Required, hot-swappable |  |  | Enter Responses Here!
Battery Operating Time, hr | At least 1 hour |  |  | Enter Responses Here!
Battery Recharging Time, hr | Bidder to specify |  |  | Enter Responses Here!
Battery Indicator | Required |  |  | Enter Responses Here!
Backup Battery | Required |  |  | Enter Responses Here!
UPS Requirement | Bidder to specify |  |  | Enter Responses Here!
Estimated Operation Time, minutes | Bidder to specify |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Hospital Information System (HIS) | Preferred; e.g: HIS, RIS, EMR, etc |  |  | Enter Responses Here! | 2
Peripheral Device Interface | Preferred |  |  | Enter Responses Here! | 2
Ability to be Networked to Physiologic Patient Monitor | Preferred; Bidder to specify compatible brand/model |  |  | Enter Responses Here! | 1
Ability to be Networked to Central Station Monitor | Preferred; Bidder to specify compatible brand/model |  |  | Enter Responses Here! | 1
Interface with Nurse Call System | Required |  |  | Enter Responses Here!
Remote Annunciator for Alarms | Preferred; Bidder to specify compatible brand/model |  |  | Enter Responses Here! | 1
Ability for Centralised Remote Monitoring of Ventilator Alarms | Preferred; Bidder to specify compatible brand/model |  |  | Enter Responses Here! | 1
Type of Networking/Connectivity Interface | Wireless or LAN-Port, or both Bidder to specify |  |  | Enter Responses Here!
Data Port | Data port must be activated (e.g: RS232, USB) to allow download of data logs from device and service |  |  | Enter Responses Here!
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware update and upgrade at no cost |  |  | Enter Responses Here!
Remote Support Capability | Preferred; Off site diagnosis of technical issues and remote fixes. |  |  | Enter Responses Here! | 3
Cybersecurity Considerations
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | Preferred |  |  | Enter Responses Here! | 2
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here! | 1
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Does the equipment include any AI Capability? | Bidder to specify AI Capabilities. |  |  | Enter Responses Here! | 3
What is the AI capability intended to assist with? (e.g., workflow automation, diagnosis assistance, patient scheduling, other) | Bidder to specify |  |  | Enter Responses Here!
Does the AI Capability have High or Low impact on Patient Care? | Bidder to specify |  |  | Enter Responses Here!
Is this equipment that incorporates AI capabilities generally classified as high-risk or low-risk? | Bidder to specify |  |  | Enter Responses Here!
Are the AI capabilities embedded in the device or require of third party applications? | Bidder to specify |  |  | Enter Responses Here!
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Temperature Requirements | The areas temperature ranges from 22-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed. |  |  | Enter Responses Here!
Humidity Requirements | The area's relative humidity ranges between 70% and 90%. Equipment must be operational in that range. If it is not please specify the range needed. |  |  | Enter Responses Here!
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 30
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Ventilator (Neonate - NICU)
Include description and pricing for all accessories (e.g., humidifier, nebulizer, tubings, etc.) required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of Ventilator (Neonate - NICU)
Include description and pricing for all consumable items required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Ventilator (Neonate - NICU)
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!