## CT Scan (256 Slice)
CT Scan (256 slice)
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 15956 Scanning Systems, Computed Tomography, Axial, Full-Body (256-slices) Total Quantity: xx | Equipment Location (Region): Quantity SIHREN a. Indonesia (Nationwide): xx units | Description: Axial computed tomography scanning systems designed to generate an image of the tissue density in any transverse plane of the whole human body. It is used for a wide variety of diagnostic procedures, including spine and head injuries, lesions, and abdominal and pelvic malignancies; to examine the cerebral ventricles, the chest wall, and the large blood vessels; and to assess musculoskeletal degeneration.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Noninvasive radiographic technique that involves the reconstruction of a tomographic plane of the body (a slice) from a large number of collected x-ray absorption measurements taken during a scan around the body's periphery. | Wide variety of imaging exams, including spine and head, gastrointestinal, and vascular |  |  | Enter Responses Here!
Coronary artery calcification scoring | Required |  |  | Enter Responses Here!
Auto vessel mapping | Required |  |  | Enter Responses Here!
Cardiac Quantification | Required |  |  | Enter Responses Here!
Ventricular output | Required |  |  | Enter Responses Here!
Lung nodule assisted reading | Required |  |  | Enter Responses Here!
Lung nodule CAD | Required |  |  | Enter Responses Here!
Lung function analysis | Required |  |  | Enter Responses Here!
Vessel analysis (noncardiac) | Required |  |  | Enter Responses Here!
Brain perfusion | Required |  |  | Enter Responses Here!
Z-axis coverage for brain perfusion | Required |  |  | Enter Responses Here!
Auto bone removal | Required |  |  | Enter Responses Here!
Body perfusion | Required |  |  | Enter Responses Here!
Virtual colonoscopy assisted reading | Required |  |  | Enter Responses Here!
Virtual colonoscopy CAD | Required |  |  | Enter Responses Here!
Myocardial evaluation | Required |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. CT Scan (256-slices)
Type | 256 slices, or equivalent |  |  | Enter Responses Here!
Number of Channel | Bidder to Specify |  |  | Enter Responses Here!
Gantry
Gantry Tilt (Mechanical or Digital), ° | At least ±30 degrees |  |  | Enter Responses Here!
Gantry Aperture (opening), cm | 70 cm and above |  |  | Enter Responses Here!
Scan Localizer | Laser |  |  | Enter Responses Here!
X-Ray Tube
Heat Storage, MHU | 7 MHU |  |  | Enter Responses Here!
Heat Dissipation Rate, kHU/min | At least 1,300 kHU/min |  |  | Enter Responses Here!
Heat Dissipation Rate, kHU/min | 2,000 kHU/min and above |  |  | Enter Responses Here!
Tube Cooling | Bidder to specify |  |  | Enter Responses Here!
Tube Focal Spots, mm | Bidder to specify |  |  | Enter Responses Here!
Expected tube life, scan sec (and calendar) | At least 200,000 scan sec |  |  | Enter Responses Here!
Expected tube life, scan sec (and calendar) | At least 300,000 scan sec |  |  | Enter Responses Here!
Max mA for smallest tube spot | At least 300 mA |  |  | Enter Responses Here!
Scan Time at Maximum mA, sec | ≥ 2 sec |  |  | Enter Responses Here!
X-Ray Generator
kW Output | At least 70 kW |  |  | Enter Responses Here!
kW Output | 100 kW and above |  |  | Enter Responses Here!
kVp Range | 70 or 80 kVp, to 140 kVp |  |  | Enter Responses Here!
mA Range | At least 20mA to 500 mA |  |  | Enter Responses Here!
Detector
Field of view, cm | ≥ 50 cm |  |  | Enter Responses Here!
Field of view, cm | ≥ 70 cm |  |  | Enter Responses Here!
Total Detector Width, z-axis, mm | At least 38mm Wider coverage provide faster scan time without compromising image quality. |  |  | Enter Responses Here!
Reconstructed Slice Width Options, mm | 0.625 (0.4) mm to 10mm. Smaller slices provide greater spatial resolution but require a larger radiation dose. |  |  | Enter Responses Here!
Optional Minimum Slice Width, mm | 0.5 mm or smaller |  |  | Enter Responses Here!
Standard Rotation Times 360°, sec | 0.4 sec to 2 sec. Shorter rotation times provide shorter scan times, better temporal resolution (less bluring in the images) |  |  | Enter Responses Here!
Standard Rotation Times 360°, sec | < 0.4 sec, or faster |  |  | Enter Responses Here!
Patient Table
Vertical movement range, cm | Lowered to at least 55 cm, or lower, from the floor |  |  | Enter Responses Here!
Minimum Longitudinal Range of Movement, cm | At least 150 cm |  |  | Enter Responses Here!
Available Longitudinal Range of Movement, cm | At least 200 cm |  |  | Enter Responses Here!
Minimum Scannable Range, cm | At least 150 cm |  |  | Enter Responses Here!
Available Scannable Range, cm | At least 200 cm |  |  | Enter Responses Here!
Maximum Load Capacity, kg | At least 200 kg |  |  | Enter Responses Here!
Maximum Load Capacity, kg | At least 220 kg |  |  | Enter Responses Here!
Performance
High-Contrast Spatial Resolution, mm | Bidder to specify |  |  | Enter Responses Here!
Modulation Transfer Function (MTF) Kernel | Bidder to specify |  |  | Enter Responses Here!
i. 0% MTF, lp/cm | ≥ 15 lp/cm |  |  | Enter Responses Here!
ii. 10% MTF, lp/cm | ≥ 10 lp/cm |  |  | Enter Responses Here!
iii. 50% MTF, lp/cm | ≥ 5 lp/cm |  |  | Enter Responses Here!
Low-Contrast Resolution, mm at % at ≤20 mGy (2 rads) | At least 5mm or lower, at 0.3% |  |  | Enter Responses Here!
Low-Contrast Resolution, mm at % at ≤20 mGy (2 rads) | 3mm or lower, at 0.3% |  |  | Enter Responses Here!
Noise, % at ≤25 mGy (2.5 rads) | Bidder to specify |  |  | Enter Responses Here!
Radiation Dose
Dose-Modulation technique | Required, to reduce radiation dose to patient |  |  | Enter Responses Here!
Pediatric-Specific Dose Control | Preferred |  |  | Enter Responses Here!
Prospective ECG Gating | Required |  |  | Enter Responses Here!
Retrospective ECG Editing | Required |  |  | Enter Responses Here!
Iterative Image Reconstruction | Preferred |  |  | Enter Responses Here!
Sliding Collimation (Overbeaming Reduction) | Required |  |  | Enter Responses Here!
Axial Cardiac | Required; Low-dose cardiac (axial acquisition) |  |  | Enter Responses Here!
Arrhythmia Correction | Required |  |  | Enter Responses Here!
Max Heart Rate, bpm | Bidder to specify |  |  | Enter Responses Here!
Clinical Software Applications
Shall come with the following softwares and protocols, but not limited to: | Image analysis software for different imaging. Required |  |  | Enter Responses Here!
Head and neck protocols | Required |  |  | Enter Responses Here!
Musculoskeletal (MSK) protocols | Required |  |  | Enter Responses Here!
Thoracic protocols | Required |  |  | Enter Responses Here!
Oncology protocols | Required |  |  | Enter Responses Here!
Perfusion protocols | Required |  |  | Enter Responses Here!
Abdominal protocols | Required |  |  | Enter Responses Here!
Multiplanar reconstruction (MPR) | Required |  |  | Enter Responses Here!
Curved MPR | Required |  |  | Enter Responses Here!
Maximum Intensity Projection (MIP) | Required |  |  | Enter Responses Here!
Minimum Intensity Projection (MIP) | Required |  |  | Enter Responses Here!
Tissue transparent view | Required |  |  | Enter Responses Here!
CT angiography (CTA) | Required |  |  | Enter Responses Here!
Volume Rendering Technique (VRT) | Required |  |  | Enter Responses Here!
Spiral scan noise reduction software | Required |  |  | Enter Responses Here!
CTA digital subtract angiography software | Required |  |  | Enter Responses Here!
3D volume measurement software | Required |  |  | Enter Responses Here!
Lung enhancement software | Required |  |  | Enter Responses Here!
Low dose lung scan | Required |  |  | Enter Responses Here!
CT cine | Required |  |  | Enter Responses Here!
Dynamic scan CT time-intensity line | Required |  |  | Enter Responses Here!
Beam-harden artifact correction software | Required |  |  | Enter Responses Here!
Parallel reconstruction | Required |  |  | Enter Responses Here!
Pediatric scan | Required |  |  | Enter Responses Here!
Metal artifacts reduction | Required |  |  | Enter Responses Here!
Neuro DSA | Required |  |  | Enter Responses Here!
Aneurysm segmentation | Required |  |  | Enter Responses Here!
Lung tumor, nodule analysis software. | Required |  |  | Enter Responses Here!
Perfusion imaging software. Provision of perfusion parameters, such as BV, BF, PPS & etc. | Required |  |  | Enter Responses Here!
Virtual endoscopy software | Required |  |  | Enter Responses Here!
Others, if any | Bidder to specify |  |  | Enter Responses Here!
Advanced Image Acquisition
Perfusion Imaging | Required |  |  | Enter Responses Here!
Cardiac | Required |  |  | Enter Responses Here!
Extended Coverage for 4-D Imaging | Preferred |  |  | Enter Responses Here!
Dual-Energy Acquisition | Preferred |  |  | Enter Responses Here!
Others, if any | Bidder to specify |  |  | Enter Responses Here!
Image Reconstruction | Required, including 3-D reconstruction feature. |  |  | Enter Responses Here!
Reconstruction FOVs, cm | At least 50 cm |  |  | Enter Responses Here!
Reconstruction Matrices | 512 x 512 |  |  | Enter Responses Here!
Maximum Reconstruction Rate, (512 x 512), image/sec | At least 20 images/sec |  |  | Enter Responses Here!
Per slice, sec | ≤ 0.2 sec |  |  | Enter Responses Here!
Real-time partial image reconstruction | Required |  |  | Enter Responses Here!
Image Processing
Recommended Post-Processing Workstation | Bidder to specify |  |  | Enter Responses Here!
Remote Access to Raw Image Data | Required |  |  | Enter Responses Here!
Remote Access to Clinical Applications | Required |  |  | Enter Responses Here!
DICOM 3D Image Export | Required |  |  | Enter Responses Here!
Control Console | To allow images to be stored, recalled, and manipulated |  |  | Enter Responses Here!
Number of Display Monitors | At least 2 monitor |  |  | Enter Responses Here!
Display Monitor Size | ≥ 19" |  |  | Enter Responses Here!
Display Monitor Size | ≥ 21" |  |  | Enter Responses Here!
Display Monitor Resolution | At least 1 Megapixels |  |  | Enter Responses Here!
Display Monitor Resolution | 2 Megapixels and above |  |  | Enter Responses Here!
Display Type | Colour LCD |  |  | Enter Responses Here!
Display Type | Colour LCD with touchscreen |  |  | Enter Responses Here!
Standard User Interface | Bidder to specify |  |  | Enter Responses Here!
Operating System | Bidder to specify |  |  | Enter Responses Here!
Number of Stored Scanning Protocols | Bidder to specify |  |  | Enter Responses Here!
Local Hard Disk Capacity | Bidder to specify |  |  | Enter Responses Here!
Viewing Workstation | For observation of CT images |  |  | Enter Responses Here!
Number of Display Monitors | At least 2 monitors |  |  | Enter Responses Here!
Number of Display Monitors | 3 or more monitors |  |  | Enter Responses Here!
Display Monitor Size | ≥ 21" |  |  | Enter Responses Here!
Display Monitor Size | ≥ 24" |  |  | Enter Responses Here!
Display Monitor Resolution | At least 2 Megapixels |  |  | Enter Responses Here!
Display Monitor Resolution | 3 Megapixels and above |  |  | Enter Responses Here!
Standard User Interface | Bidder to specify |  |  | Enter Responses Here!
Number of Laser Printers | ≥ 1 |  |  | Enter Responses Here!
Storage Server | Images to be stored in the system for later retrieval and review/ manipulation, hard copies can be printed from digital storage. |  |  | Enter Responses Here!
Operating System | Bidder to specify |  |  | Enter Responses Here!
Image Storage Capacity, Local Hard Drive, TB | Bidder to specify |  |  | Enter Responses Here!
Contrast Injection Integration | Required |  |  | Enter Responses Here!
b. Contrast Media Injectors
Type | Contrast media delivery system for CT Scanner |  |  | Enter Responses Here!
Injector Head Mounting Configuration | Bidder to specify |  |  | Enter Responses Here!
Arm Switch Control | Required, to prevent the accidental triggering of an injection |  |  | Enter Responses Here!
Fluid Heating Feature | Required, to maintain the contrast medium at or near body temperature |  |  | Enter Responses Here!
Temperature Control | Thermostatically controlled |  |  | Enter Responses Here!
Temperature Range, ⁰C | Bidder to specify |  |  | Enter Responses Here!
Syringe Injection/Delivery System | Automatic, Electromechanically operated |  |  | Enter Responses Here!
Syringe Type | Dual-syringe type |  |  | Enter Responses Here!
Disposable Syringe | Disposable type to prevent the spread of infection to subsequent patients. |  |  | Enter Responses Here!
Syringe Transparency | Required, to allow air bubbles to be seen more easily and purged before the injection |  |  | Enter Responses Here!
Contrast Syringe Capacity, mL | Up to 200 mL. Bidder to specify range. |  |  | Enter Responses Here!
Saline Syringe Capacity, mL | Bidder to specify range. |  |  | Enter Responses Here!
Flow Controlled Feature | Required, to control delivery of contrast medium |  |  | Enter Responses Here!
Flow Rate Range, mL/sec | Adjustable from 0.1mL/sec to 10mL/sec. |  |  | Enter Responses Here!
Pressure Controlled Feature | Required, to control the maximum pressure that can be generated in the syringe |  |  | Enter Responses Here!
Delivery Pressure Range, psi | At least 50 to 325 psi. |  |  | Enter Responses Here!
Increment, psi | In increment of 10 psi or lower. |  |  | Enter Responses Here!
Rise Time Feature | Adjustable, the amount of time required to achieve proper pressure/flow after initiating an injection to reduce the incidence of catheter dislocation and vascular trauma. |  |  | Enter Responses Here!
Increment, sec | Bidder to specify |  |  | Enter Responses Here!
Adjustable Volume Stop (Backup Stop) | Required, to prevent the excess delivery of contrast medium |  |  | Enter Responses Here!
Mechanism | Electronically or Mechanically. |  |  | Enter Responses Here!
Increment, mL | In increment of 1.0 mL or lower. |  |  | Enter Responses Here!
Tubing Saline Flush | Required, automatically switching from contrast medium to saline once contrast optimization has been achieved, to reduce the amount of medium that the patient receives. |  |  | Enter Responses Here!
Alerts System | Required, audible and visual |  |  | Enter Responses Here!
Air Detected and/or Purge Required | Alarm triggered |  |  | Enter Responses Here!
Low or Empty Syringe Volume | Alarm triggered |  |  | Enter Responses Here!
High Flow and/or Volume Injected | Alarm triggered |  |  | Enter Responses Here!
Overpressure | Alarm triggered |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
X-Ray Generator Synchronization | Required |  |  | Enter Responses Here!
User Interface | Remote controlled console |  |  | Enter Responses Here!
Display Monitor | Colour touchscreen |  |  | Enter Responses Here!
Screen Size, inches | Bidder to specify |  |  | Enter Responses Here!
Protocols Storage | Programmable |  |  | Enter Responses Here!
Maximum Numbers of Protocols Stored | Bidder to specify |  |  | Enter Responses Here!
Maximum Numbers of Phases | Bidder to specify |  |  | Enter Responses Here!
Application Software | Bidder to provide and specify all software required for intended procedures |  |  | Enter Responses Here!
Air Detector System | Required; Air-in-line sensor to alert the user and stop the injection if air is detected |  |  | Enter Responses Here!
Extravasation (Leakage) Detection System | Required; To prevent the leaking of contrast media from the veins into surrounding tissue |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. CT Scan (256-slices)
a) Phantoms, Computed Tomography | Bidder to supply complete set for performance evaluation, quality assurance and quality control. Bidder to supply the following according to IAEA standards, but not limited to: 1. CTDI Phantom Set 2. Image Quality Phantom |  |  | Enter Responses Here!
b) Standard Sofware Modules/Functionality | Bidder to specify |  |  | Enter Responses Here!
c) Additional Software Modules/Functionality | Bidder to specify |  |  | Enter Responses Here!
d) Radiation Protective Equipment | Bidder to supply the following: 1. Lead Apron (2 pcs per unit) 2. Thyroid Shield (2 pcs per unit) 3. Pb Glasses (2 pcs per unit) |  |  | Enter Responses Here!
e) Radiation Monitoring Devices | Bidder to supply the following: 1. Surveymeter (2 pcs per unit) 2. Analog/ Digital Pocket Dosimeter (2 pcs per unit) 3. Film Badge/ TLD (2 pcs per unit) |  |  | Enter Responses Here!
Others, if any | Bidder to specify |  |  | Enter Responses Here!
b. Contrast Media Injectors
a) Adapter for Syringes | 1 set per unit |  |  | Enter Responses Here!
b) Contrast Warmer Cabinet | Required |  |  | Enter Responses Here!
Others, if any | Bidder to specify |  |  | Enter Responses Here!
START-UP CONSUMABLES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. CT Scan (256-slices)
N/A | N/A
b. Contrast Media Injectors
Disposable Syringes and Tubings Kit | Bidder to supply quantity for 100 typical patient procedures |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dimension, H x W x D, cm
a. CT Scan (64-slices) | Bidder to specify |  |  | Enter Responses Here!
b. Contrast Media Injectors | Bidder to specify |  |  | Enter Responses Here!
Weight, kg
a. CT Scan (64-slices) | Bidder to specify |  |  | Enter Responses Here!
b. Contrast Media Injectors | Bidder to specify |  |  | Enter Responses Here!
MECHANICAL & ELECTRICAL (M&E) REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. CT Scan (256-slices)
Electrical Requirement, V | 3-phase 440 V AC, 50Hz |  |  | Enter Responses Here!
Power Supply | Compatible isolator with equipment, and in compliance with Indonesia requirement |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Scanning, kW | Bidder to specify |  |  | Enter Responses Here!
Standby, kW | Bidder to specify |  |  | Enter Responses Here!
Not in use, kW | Bidder to specify |  |  | Enter Responses Here!
UPS Requirement | Required; for CT Scanner and Workstation |  |  | Enter Responses Here!
Estimated Operation Time, minutes | At least 30 minutes |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here!
b. Contrast Media Injectors
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here!
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | Required; e.g: HIS, RIS, EMR, etc |  |  | Enter Responses Here!
Compatibility with DICOM (Digital Imaging and Communications in Medicine) | Required; DICOM 3.0, PACS Connectivity, DICOM encryption support |  |  | Enter Responses Here!
DICOM Modality Worklist | Preferred |  |  | Enter Responses Here!
Type of Networking/Connectivity Interface | Wireless or LAN-Port, or both Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
Data Port | Data port must be activated (e.g: RS232, USB) to allow download of data logs from device and service |  |  | Enter Responses Here!
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware updates and upgrades at no cost |  |  | Enter Responses Here!
Remote Support Capability | Preferred; Off site diagnosis of technical issues and remote fixes. |  |  | Enter Responses Here!
Cybersecurity Considerations |  |  |  | Enter Responses Here!
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | Preferred |  |  | Enter Responses Here!
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here!
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Artficial Intelligence capabilities | Required; Bidder to list all AI Capabilities including software packages (e.g: image registration/ optimisation/ analysis/ segmentation/ lesion detection and localisation, quantitative analysis, automated reporting/ predictive modelling, calculations, etc.) |  |  | Enter Responses Here!
AI Capabilities: Intended Use: Assist radiologists and neurologists in the analysis of non-contrast and contrast-enhanced CT scans for detecting, characterizing, and measuring acute ischemic stroke. AI Integration: Preferably, the AI capabilities should be embedded within the CT scanner software for seamless integration and workflow efficiency. | Preferred |  |  | Enter Responses Here!
Specific functionalities of the AI application may include:
Automatic detection of ischemic lesions: The application should identify and highlight potential stroke regions within the CT scan with high accuracy. | Preferred |  |  | Enter Responses Here!
Volume measurement: The application should automatically calculate the volume of the ischemic lesion. | Preferred |  |  | Enter Responses Here!
Stroke subtype classification: The application should predict the subtype of stroke (e.g., large vessel occlusion, small vessel occlusion) with high accuracy. | Preferred |  |  | Enter Responses Here!
Collaterals assessment: If possible, the application should analyze collateral blood flow to the ischemic region to inform treatment decisions. | Preferred |  |  | Enter Responses Here!
Integration with clinical scoring systems: The application should integrate with established stroke scoring systems (e.g., ASPECTS, NIHSS) to facilitate clinical decision-making. | Preferred |  |  | Enter Responses Here!
What is the AI capability intended to assist with? (e.g., workflow automation, diagnosis assistance, patient scheduling, other) | Bidder to specify |  |  | Enter Responses Here!
Does the AI Capability have High or Low impact on Patient Care? | Bidder to specify |  |  | Enter Responses Here!
Is this equipment that incorporates AI capabilities generally classified as high-risk or low-risk? | Bidder to specify |  |  | Enter Responses Here!
Are the AI capabilities embedded in the device or require of third party applications? | Bidder to specify |  |  | Enter Responses Here!
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Telemedicine Capabilities | Bidder to list all telemedicine capabilities |  |  | Enter Responses Here!
Asynchronous telemedicine capabilities. Store-and-forward technology is asynchronous in that it allows for the images or data to be interpreted at a later time and eliminates the need for the patient and physician to be in the same room or venue. (i.e., through store and forward, in which patient information, such as an image or data, "stored" in the originating facility's (i.e., where the patient is) computer system is sent securely ("forwarded") to the telehealth provider at the distant site (i.e., where the provider is) for review.) | Bidder to specify |  |  | Enter Responses Here!
Remote patient monitoring capabilities. Remote patient monitoring, in which vital sign information (e.g., blood pressure) and other data (e.g., blood glucose levels) are collected by monitoring devices, usually at home, and transferred to the patient's physician. This technology is used often in the care of patients with chronic conditions. | Bidder to specify |  |  | Enter Responses Here!
Mobile health (mHealth) integration capabilities. Mobile health (mHealth), in which healthcare and public health information are communicated through mobile devices. The information may include general educational information, targeted texts, and notifications about disease outbreaks. | Bidder to specify |  |  | Enter Responses Here!
Other telemdicine capabilities. | Bidder to specify |  |  | Enter Responses Here!
Is telemedicine offered through a 3rd party vendor? | Bidder to specify |  |  | Enter Responses Here!
Video Capabilities (if available)
Are video capabilities needed? If so, specify if they are included in the solution or offered through 3rd party add-on? | Bidder to specify |  |  | Enter Responses Here!
If video capabilities are available, specify the resolution | Bidder to specify |  |  | Enter Responses Here!
Is night vision available? | Bidder to specify |  |  | Enter Responses Here!
Does the system offer pan/tilt features? | Bidder to specify |  |  | Enter Responses Here!
Is video zoom available? | Bidder to specify |  |  | Enter Responses Here!
Audio capabilities (if available)
Is audio functionality available? | Bidder to specify |  |  | Enter Responses Here!
Specify audio features | Bidder to specify |  |  | Enter Responses Here!
Is echo cancellation available? | Bidder to specify |  |  | Enter Responses Here!
Is noise cancellation available? | Bidder to specify |  |  | Enter Responses Here!
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Equipment Room Readiness Requirement | To supply, install and complete commission of the room including radiation protection (in accordance to BAPETEN and PERMENKES guidelines) |  |  | Enter Responses Here!
Audio/Video Communication System | Bidder to provide 2-ways system. To be integrated with Control Console
Pre-Installation Requirements | All electrical power supply and installation to technical, control and equipment room
Room Radiation Protection | Bidder to supply and install equivalent lead thickness
CT Scanner Unit Placement | Bidder to provide floor plan drawing
Air conditioning requirements | To supply, install and complete commissioning of required air conditioning in equipment room and technical room
Fire Alarm and Protection System | Bidder to supply and install
Control Room or Workstation | Bidder to supply and install
Technical Room | Bidder to supply and install
Temperature Requirements | The areas temperature ranges from 20-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed.
Humidity Requirements | The area's relative humidity ranges between 70% and 90%. Equipment must be operational in that range. If it is not please specify the range needed.
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. Bidder to specify sessions for various user groups (e.g. Operators, Oncologists, Medical Physicists, Biomedical Engineers,etc.) |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extension of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of CT Scan (256 slice)
Include description and pricing for all accessories offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of CT Scan (256 slice)
Include description and pricing for all consumable items required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of CT Scan (256 slice)
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!