## Mammography
Mammography
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS 18432 Radiographic Systems, Digital, Mammographic Total Quantity: 354 | Equipment Location (Region): Quantity SIHREN: a. Indonesia (Nationwide): 354 units | Description: Digital radiographic systems designed for radiography of the breast (i.e., mammography). These systems usually include a high-frequency, constant-potential x-ray generator; an automatic exposure control device; filters; a collimator; patient positioning features (e.g., compression devices); and a digital imaging system. Typically, the digital imaging system consists of a phosphor imaging plate used to obtain a latent image, which is scanned with a laser and digitized in an analog-to-digital converter. Digital mammographic systems are specifically intended for imaging the breast; they are used for screening and/or assessment of breast lesions.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIRED SPECIFICATIONS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dedicated for 3D breast scanning | Screening and/or assessment of breast lesions, detection and diagnosis of breast cancer and to evaluate palpable masses and nonpalpable breast lesions. |  |  | Enter Responses Here!
For Full Field Digital Mammography (FFDM) | Use of low-energy radiation to distinguish between normal and pathologic soft tissue, to produce high-resolution images, and to visualize the subtle signs of early breast cancer. |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIRED SPECIFICATIONS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Type | Mammography System |  |  | Enter Responses Here!
X-Ray Generator | High frequency type |  |  | Enter Responses Here!
kV Range | Within 20 to 50 kV |  |  | Enter Responses Here!
Minimum kV increments | 0.1 to 1 kV |  |  | Enter Responses Here!
mAs Range | Within 0.1 - 650 mAs |  |  | Enter Responses Here!
mA Range | Within 20 - 200 |  |  | Enter Responses Here!
Time Range, sec | Within 0.01 - 8 |  |  | Enter Responses Here!
X-Ray Tube and Gantry | Rotating anode type |  |  | Enter Responses Here!
Heat Capacity, HU | ≥ 300,000 HU Indication of the maximum heating during x-ray generation that the x-ray tube anode can withstand. |  |  | Enter Responses Here!
Heat Dissipation Rate, HU/min | At least 40,000 |  |  | Enter Responses Here!
Heat Dissipation Rate, HU/min | 60,000 (HU/min) and above |  |  | Enter Responses Here!
Cooling Method | Oil-cooled and/or Fan-cooled. Bidder to specify |  |  | Enter Responses Here!
Anode Target/Filter Combinations | Bidder to specify |  |  | Enter Responses Here!
Focal Spot Size, mm | 0.1mm to 0.3 mm, to ensure that microcalcifications as small as 200 µm can be detected. |  |  | Enter Responses Here!
Digital Detector | Removable photostimulable storage phosphor detectors, or Amorphous selenium (direct) flat panels, Amorphous silicon flat panels, or Cesium Iodide (CsI) Bidder to specify |  |  | Enter Responses Here!
Detector Size, W x H, cm | At least 23cm x 26 cm, to accommodate all breast sizes with a single exposure. |  |  | Enter Responses Here!
Maximum Exposure Time, sec | Bidder to specify |  |  | Enter Responses Here!
Image Bit Depth | 14 bits |  |  | Enter Responses Here!
Image Bit Depth | 16 bits |  |  | Enter Responses Here!
Pixel Size, µm | ≤ 100 µm; |  |  | Enter Responses Here!
Pixel Size, µm | ≤ 50 µm; |  |  | Enter Responses Here!
Detective Quantum Efficiency (DQE) at 1 lp/mm | ≥50% @ 28 kV Bidder to specify |  |  | Enter Responses Here!
Detective Quantum Efficiency (DQE) at 5 lp/mm | ≥20% @ 28 kV at 5 lp/mm Bidder to specify |  |  | Enter Responses Here!
Modulation Transfer Function (MTF) at 5 lp/mm | ≥40% Bidder to specify |  |  | Enter Responses Here!
Grid System | Required, to reduce scatter |  |  | Enter Responses Here!
Grid Ratio | Bidder to specify |  |  | Enter Responses Here!
Positioning Assembly | Automatic or Manual collimation. |  |  | Enter Responses Here!
Collimation Size | At least 18cm x 23 cm; and at least 24cm x 29 cm |  |  | Enter Responses Here!
Movement Locking Mechanism | Required Bidder to specify type |  |  | Enter Responses Here!
Assembly Rotation Movement, ° | Bidder to specify |  |  | Enter Responses Here!
Vertical Movement Range, cm | At least 80 - 130 cm, a wider range is acceptable |  |  | Enter Responses Here!
Vertical Movement Range, cm | At least 70 - 140 cm, a wider range is acceptable |  |  | Enter Responses Here!
Source-to-Image Distance (SID) Movement, cm | ≥65 cm, to clearly image the smallest microcalcifications |  |  | Enter Responses Here!
Scale Guide | Required, for distance and pressure parameters. To indicate the amount of compression applied to the breast during scanning |  |  | Enter Responses Here!
Height Adjustable | Required, Low height sufficient for scanning in sitting and lateral position. |  |  | Enter Responses Here!
Compression System | Manual/Motorized, Automatic, Fine Adjustment |  |  | Enter Responses Here!
Stop or Release Button | Required, manually or automatically released at a preset pressure or during emergency cases |  |  | Enter Responses Here!
Compression Force, N | Maximum 200N |  |  | Enter Responses Here!
Magnification Device | Capability to increase object-to-image distance (OID) to obtain more diagnostic information and to reduce the need for surgical biopsy. |  |  | Enter Responses Here!
Automatic Exposure Control (AEC) Detector | Required |  |  | Enter Responses Here!
Type | Automatic or Semi-automatic or Manual. |  |  | Enter Responses Here!
Parameters Controlled | kV, mAs, and Anode/Filter |  |  | Enter Responses Here!
Minimum Radiation Output, mR/sec @ 28 kVp | At least 200 mR/sec |  |  | Enter Responses Here!
Radiation Output, mR/sec @ 28 kVp | 500 mR/sec or more |  |  | Enter Responses Here!
Radiation Shield | Required. Bidder to specify shield specification |  |  | Enter Responses Here!
Dimension, L x W, cm (in) | Bidder to specify |  |  | Enter Responses Here!
Thickness | Bidder to specify |  |  | Enter Responses Here!
Bucky | Required; at least size 24 x 29 cm Bidder to specify other sizes if applicable |  |  | Enter Responses Here!
Mammography Chair | Mobile, electrical operated |  |  | Enter Responses Here!
Number of sections | At least 3 sections |  |  | Enter Responses Here!
Height Adjustable | Required |  |  | Enter Responses Here!
Height Adjustment Range, cm | Bidder to specify |  |  | Enter Responses Here!
Adjustable Back Support | Required |  |  | Enter Responses Here!
Height Adjustable Head Support | Required |  |  | Enter Responses Here!
Drop-down Shoulder Sections | Required |  |  | Enter Responses Here!
Foldable Arm Rests | Required |  |  | Enter Responses Here!
Height Lock | Safety switch to lock height adjustment during procedure |  |  | Enter Responses Here!
Lockable Castors | Required |  |  | Enter Responses Here!
Maximum Patient Weight | ≥200kg |  |  | Enter Responses Here!
Chair Upholstery | Electrical conductive |  |  | Enter Responses Here!
Display Monitors | Required |  |  | Enter Responses Here!
Type | LCD or LED display |  |  | Enter Responses Here!
Screen Size, in | ≥ 19 inches |  |  | Enter Responses Here!
Display Resolution, MP | At least 3 MP |  |  | Enter Responses Here!
Display Resolution, MP | 5 MP or more |  |  | Enter Responses Here!
Compatible to display 3D images | Required |  |  | Enter Responses Here!
Operating System | Bidder to specify |  |  | Enter Responses Here!
Number of Stored Scanning Protocols | Bidder to specify |  |  | Enter Responses Here!
Local Hard Disk Capacity | At least 1 TB |  |  | Enter Responses Here!
Storage Memory, RAM | Bidder to specify |  |  | Enter Responses Here!
Data Back-up System | Preferred;. Bidder to state additional storage capacity in GBs |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) Phantoms, Mammographic | Complete set for performance evaluation, quality assurance and quality control. Bidder to supply the following according to IAEA standards, but not limited to: 1. Phantom Mammography Image Quality 2. Phantom PMMA Thickness Variations 3. Resolution Pattern Test Tool 4. Radiographic Ruler Mammography Tool |  |  | Enter Responses Here!
b) Compression Paddles | Complete set for screening and diagnostic. Bidder to supply at least 3 appropriate sizes of paddles, to accommodate various breast size, shape and comfort level. |  |  | Enter Responses Here!
c) Radiation Protective Equipment | Bidder to supply the following: 1. Lead Apron (2 pcs per unit) 2. Thyroid Shield (2 pcs per unit) 3. Face Shield (2 pcs per unit) 3. Pb Glasses (2 pcs per unit) 4. Pb Gloves (2 pcs per unit) |  |  | Enter Responses Here!
d) Radiation Monitoring Devices | Bidder to supply the following: 1. Surveymeter (2 pcs per unit) 2. Analog/ Digital Pocket Dosimeter (2 pcs per unit) 3. Film Badge/ TLD (2 pcs per unit) |  |  | Enter Responses Here!
d) Others, if any | Bidder to list and supply |  |  | Enter Responses Here!
START-UP CONSUMABLES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Mammography System
Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Lowest Height Adjustment, H, cm | Bidder to specify |  |  | Enter Responses Here!
Weight, kg | Bidder to specify |  |  | Enter Responses Here!
Mammography Chair
Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
ELECTRICAL REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Mammography System
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!+E113
Power Supply | Compatible isolator with equipment, and in compliance with Indonesia requirement |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Scanning, kW | Bidder to specify |  |  | Enter Responses Here!
Standby, kW | Bidder to specify |  |  | Enter Responses Here!
Not in use, kW | Bidder to specify |  |  | Enter Responses Here!
Battery Back-up Power (UPS) | Rechargeable |  |  | Enter Responses Here!
Battery Operating Time, hr | Bidder to specify |  |  | Enter Responses Here!
Battery Recharging Time, hr | Bidder to specify |  |  | Enter Responses Here!
Mammography Chair
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here!
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | Required, e.g: HIS, RIS, EMR, etc |  |  | Enter Responses Here!
Compatibility with DICOM (Digital Imaging and Communications in Medicine) | Required; DICOM 3.0, PACS Connectivity, DICOM encryption support |  |  | Enter Responses Here!
Type of Networking/Connectivity Interface | Wireless or LAN-Port, or both Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
Data Port | Data port must be activated, e.g: RS232 Bidder to ensure all output data will be transferable or integrated into HIS. |  |  | Enter Responses Here!
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware updates and upgrades at no cost |  |  | Enter Responses Here!
Remote Support Capability | Preferred; Off site diagnosis of technical issues and remote fixes. |  |  | Enter Responses Here!
Cybersecurity Considerations
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | Required. |  |  | Enter Responses Here!
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here!
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Does the equipment include any AI Capability? | Bidder to specify AI Capabilities. |  |  | Enter Responses Here!
What is the AI capability intended to assist with? (e.g., workflow automation, diagnosis assistance, patient scheduling, other) | Bidder to Specify |  |  | Enter Responses Here!
Does the AI Capability have High or Low impact on Patient Care? | Bidder to Specify |  |  | Enter Responses Here!
Is this equipment that incorporates AI capabilities generally classified as high-risk or low-risk? | Bidder to specify |  |  | Enter Responses Here!
Are the AI capabilities embedded in the device or require of third party applications? | Bidder to specify |  |  | Enter Responses Here!
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Does the equipment have telemedicine capability or can it be used in telemedicine applications? | Bidder to specify telemedicine capabilities |  |  | Enter Responses Here!
Asynchronous telemedicine capabilities. Store-and-forward technology is asynchronous in that it allows for the images or data to be interpreted at a later time and eliminates the need for the patient and physician to be in the same room or venue. (i.e., through store and forward, in which patient information, such as an image or data, "stored" in the originating facility's (i.e., where the patient is) computer system is sent securely ("forwarded") to the telehealth provider at the distant site (i.e., where the provider is) for review.) | Bidder to specify |  |  | Enter Responses Here!
Remote patient monitoring capabilities. Remote patient monitoring, in which vital sign information (e.g., blood pressure) and other data (e.g., blood glucose levels) are collected by monitoring devices, usually at home, and transferred to the patient's physician. This technology is used often in the care of patients with chronic conditions. | Bidder to specify |  |  | Enter Responses Here!
Mobile health (mHealth) integration capabilities. Mobile health (mHealth), in which healthcare and public health information are communicated through mobile devices. The information may include general educational information, targeted texts, and notifications about disease outbreaks. | Bidder to specify |  |  | Enter Responses Here!
Other telemedicine capabilities. | Bidder to specify |  |  | Enter Responses Here!
Is telemedicine offered through a 3rd party vendor? | Bidder to specify |  |  | Enter Responses Here!
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENT | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Equipment Room Readiness Requirement | To supply, install and complete commission of the room including radiation protection (in accordance to BAPETEN and PERMENKES guidelines) |  |  | Enter Responses Here!
Pre-Installation Requirements | All electrical power supply and installation to technical, control and equipment room
Room Radiation Protection | Bidder to supply and install equivalent lead thickness
Mammography Unit Placement | Bidder to provide floor plan drawing
Air conditioning requirements | To supply, install and complete commissioning of required air conditioning in equipment room and technical room
Workstation | Bidder to supply and install
Temperature Requirements | The areas temperature ranges from 20-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed.
Humidity Requirements | The area's relative humidity ranges between 70% and 90%. Equipment must be operational in that range. If it is not please specify the range needed.
Other Requirements | Bidder to specify |  |  | Enter Responses Here!
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. Bidder to specify sessions for various user groups (e.g. Operators, Oncologists, Medical Physicists, Biomedical Engineers,etc.) |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extension of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Mammography
Include description and pricing for all accessories offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Mammography
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!