## Dental Chair (Complete System)
Dental Chair (Complete System) |  |  |  |  | 70 | <--- Total score
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 11165 + 10792 Dental Delivery Units + Chairs, Procedure, Dentistry Total Quantity: 1,293 | Equipment Location (Region): Quantity SOPHI a. Indonesia (Nationwide): 2060 units | Description: A complete Dental Delivery Unit system includes: a. Dental Delivery Unit b. Dental Chair c. Medical Air Compressors d. Central Vacuum Pumps Equipment combined in a single unit designed to deliver dental care. These units typically combine all the fundamental devices needed to provide dental services, such as handpieces, lights, and cuspidors; they also include the supplies needed for their operation, such as compressed air, suction, electricity, and water. Dental delivery units are frequently available as stand-alone units, but some units are attached to a dental chair. Dental delivery units are the main component of dental systems that may also include x-ray devices, cabinets, stools, sterilization units, and other ancillary equipment. + Procedure chairs designed to facilitate dental examination, treatment, and/or minor surgical procedures. These chairs are typically adjustable up to a height that allows the healthcare staff to perform procedures while standing; the chairs usually include head- and armrests, a back that may be tilted from a vertical to near-horizontal position, and rotating capabilities to facilitate examination and/or treatment. The chairs may be mechanically operated or hydraulic or electrically powered; some chairs can be programmed for several standard positions.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
a. Dental Delivery Unit | Bidder shall include all components and features below:
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
b. Dental Chair | Bidder shall include all components and features below:
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
c. Medical Air Compressors | Bidder shall include all components and features below:
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
d. Vacuum Pumps | Bidder shall include all components and features below:
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
A complete Dental Delivery Unit designed to deliver dental care. | For Dental Specialist Clinics |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. Dental Delivery Unit
i. Dental Element | Hanging outlets system |  |  | Enter Responses Here!
Handpiece Holder | At least 4 modules |  |  | Enter Responses Here!
Handpiece Holder | 5 modules or more |  |  | Enter Responses Here! | 5
High Speed Handpiece (Contra angle) | Required, bidder to specify speed available (rpm) |  |  | Enter Responses Here!
Low Speed Handpiece (Straight) | Required, bidder to specify speed available (rpm) |  |  | Enter Responses Here!
3 Way Air/Water Syringe | Required |  |  | Enter Responses Here!
Air/Water Selection | Bidder to specify |  |  | Enter Responses Here!
Hot/Cold Selection | Bidder to specify |  |  | Enter Responses Here!
Dental Ultrasonic Scaler | Preferred, for Prophylaxis, Endodontics and Peridontics |  |  | Enter Responses Here! | 2
Drive Type | Piezoelectric |  |  | Enter Responses Here! | 2
Oscillation Frequency Range, kHz | Bidder to specify |  |  | Enter Responses Here!
Power Adjustment Range, W | Bidder to specify |  |  | Enter Responses Here!
Illumination LED Light | Bidder to specify |  |  | Enter Responses Here!
Tips Available | Bidder to specify |  |  | Enter Responses Here!
Dental Resin Polymerization Light | Required, to accelerate the polymerization process of light-cured dental resin-based composites Bidder to specify brand and model |  |  | Enter Responses Here!
Compatible Materials for Polymerization | Bidder to specify |  |  | Enter Responses Here!
Operation Mode | Bidder to specify |  |  | Enter Responses Here!
Light Source | Halogen |  |  | Enter Responses Here!
Light Source | LED (preferred) |  |  | Enter Responses Here! | 3
Light Intensity, mW/cm² | Bidder to specify |  |  | Enter Responses Here!
Light Emission Wavelength Range, nm | Bidder to specify |  |  | Enter Responses Here!
Polymerization Area, Diameter, mm | Bidder to specify |  |  | Enter Responses Here!
Curing Timer | Bidder to specify |  |  | Enter Responses Here!
Tray Holder | At least 1 standard tray |  |  | Enter Responses Here!
Tray Holder | 2 or more standard tray |  |  | Enter Responses Here! | 3
Saliva Suction/Ejector | Required |  |  | Enter Responses Here!
Control Type | Touch Pad Control and Foot Switch Control |  |  | Enter Responses Here!
Controllable Parameters for Touch Pad | Bidder to specify |  |  | Enter Responses Here!
Controllable Parameters for Foot Switch | Bidder to specify |  |  | Enter Responses Here!
X-ray Viewer | Bidder to specify |  |  | Enter Responses Here!
ii. Cuspidor/Spittoon | Required |  |  | Enter Responses Here!
Ceramic Bowl, Cup Filler and Bowl Rinsing Outlets | Removable for easy cleaning |  |  | Enter Responses Here!
Cuspidor/Spittoon Positioning | Rotatable in Swivel Direction |  |  | Enter Responses Here!
Angle of Rotation, ˚ | Bidder to specify |  |  | Enter Responses Here!
Cup Filling System and Bowl Rinsing System | Automatic Stop after the Preset Time |  |  | Enter Responses Here!
Cup Filling Time, s | Programmable Bidder to list adjustable time range for cup filling |  |  | Enter Responses Here!
Bowl Rinsing Time, s | Programmable Bidder to list adjustable time range for bowl rinsing |  |  | Enter Responses Here!
Water Heater for Cup Filler | Bidder to specify |  |  | Enter Responses Here!
Control Type | Local Control and/or Touch Pad |  |  | Enter Responses Here!
Water Bottle System | Required, to supply distilled water for handpieces, cup filler and bowl rinsing outlets |  |  | Enter Responses Here!
Capacity, L | At least 1.5 L |  |  | Enter Responses Here!
Cross Contamination Prevention | Required, to prevent contaminated water from entering water supply line |  |  | Enter Responses Here!
Disinfectant System | Required, to disinfect the spray circuits with disinfectant solution. |  |  | Enter Responses Here!
iii. Dental Air Abrasion Units | Required, to remove tooth structure by using a stream of aluminium oxide particles generated from compressed air Bidder to specify brand and model |  |  | Enter Responses Here!
Configuration | Handheld handpiece |  |  | Enter Responses Here!
Air Pressure Source | Bidder to specify: Integrated Compressor / Stand-Alone Portable Compressor / Air Supply Line from Dental Delivery Unit / Others |  |  | Enter Responses Here!
Air Pressure Adjustment Range, bar | Up to 7 bar |  |  | Enter Responses Here!
Powder Flow Rate Range, gm/min | Bidder to specify |  |  | Enter Responses Here!
Handpiece Tips | Autoclavable |  |  | Enter Responses Here!
Handpiece Nozzle Diameter Available, mm | Bidder to specify |  |  | Enter Responses Here!
Handpiece Nozzle Angulation Available, ˚ | Bidder to specify |  |  | Enter Responses Here!
Abrasive Medium | Aluminum Oxide Powder 27 µm and 50 µm |  |  | Enter Responses Here!
Water Spray System | Required, to reduce powder overspray and enhance patient comfort |  |  | Enter Responses Here!
Water Pressure Range, bar | Bidder to specify |  |  | Enter Responses Here!
Control Type | Local Control / Touch Pad / Foot Control |  |  | Enter Responses Here!
iii) Dental Light | Required |  |  | Enter Responses Here!
Light Source | Halogen |  |  | Enter Responses Here!
Light Source | LED (preferred) |  |  | Enter Responses Here! | 3
Expected Lifespan, hr | Bidder to specify |  |  | Enter Responses Here!
Colour Temperature, K | ≥ 5,000K Bidder to specify |  |  | Enter Responses Here!
Light Intensity, lux | ≥ 25,000 Lux |  |  | Enter Responses Here!
Light Intensity Selection Modes | Bidder to specify |  |  | Enter Responses Here!
Light Head Positioning | 3-axis Rotation |  |  | Enter Responses Here!
Light Handle | Removable for disinfection |  |  | Enter Responses Here!
Control Type | Local Control and/or Touch Pad |  |  | Enter Responses Here!
iv) Dental Stool | Required |  |  | Enter Responses Here!
Configuration | Mobile |  |  | Enter Responses Here!
Seat | Height-adjustable and Tiltable |  |  | Enter Responses Here!
Height Adjustment Range, cm | Bidder to specify |  |  | Enter Responses Here!
Tiltable Angle, ˚ | Bidder to specify angle measured from horizontal position |  |  | Enter Responses Here!
Seat Diameter, cm | At least 38cm Bidder to specify |  |  | Enter Responses Here!
Backrest | Tiltable |  |  | Enter Responses Here!
Tiltable Angle, ˚ | Bidder to specify angle measured from vertical position |  |  | Enter Responses Here!
Castors | 5 swivel castors |  |  | Enter Responses Here!
Material of Construction
Frame | Bidder to specify |  |  | Enter Responses Here!
Seat and Back Rest | Bidder to specify |  |  | Enter Responses Here!
Upholstery | Disinfectable Materials Bidder to specify |  |  | Enter Responses Here!
Colour Options | Bidder to specify |  |  | Enter Responses Here!
b. Dental Chair
Type | Electrical operated |  |  | Enter Responses Here!
Chair Positioning | Programmable Bidder to list all available pre-programmed positions |  |  | Enter Responses Here!
Seat | Height-adjustable and Tiltable |  |  | Enter Responses Here!
Height Adjustment Range, cm | Bidder to specify |  |  | Enter Responses Here!
Tilting Angle Range, ˚ | Bidder to specify angle from horizontal position |  |  | Enter Responses Here!
Backrest | Tiltable |  |  | Enter Responses Here!
Tilting Angle Range, ˚ | Bidder to specify angle from vertical position |  |  | Enter Responses Here!
Headrest | Double-jointed/articulating adjustable |  |  | Enter Responses Here!
Armrest | Left and Right; Swivel-jointed adjustable |  |  | Enter Responses Here!
Control Type | Footswitch and/or Touch Pad Control |  |  | Enter Responses Here!
Material of Construction | Corrossion-free, Durable, Scratch-resistant Epoxy Paint Finish |  |  | Enter Responses Here!
Frame | Bidder to specify |  |  | Enter Responses Here!
Upholstery | Seamless, Soft, easy to clean and disinfect |  |  | Enter Responses Here!
Colour Options | Bidder to specify |  |  | Enter Responses Here!
c. Medical Air Compressors
Type | Oil-free, medical grade |  |  | Enter Responses Here!
Capacity | Bidder to ensure the number of air compressors provided are sufficient to support all the required dental delivery units stipulated in the tender document |  |  | Enter Responses Here!
Maximum Number of Dental Delivery Units can be operated simultaneously | Bidder to specify |  |  | Enter Responses Here!
System Horsepower, hp | Bidder to specify |  |  | Enter Responses Here!
Tank Size, L | Bidder to specify |  |  | Enter Responses Here!
Flow Rate, L/min | At least 120L/min - 140L/min at 5 bar for each dental chair Bidder to specify flow rate at specific pressure |  |  | Enter Responses Here!
Working Pressure, bar | Up to 9 bar Bidder to specify |  |  | Enter Responses Here!
Cut-off Pressure, bar | Bidder to specify |  |  | Enter Responses Here!
Pressure Gauges | Required, to monitor the air in the tank and outlet |  |  | Enter Responses Here!
Safety Valve | Required |  |  | Enter Responses Here!
Maximum Permissible Pressure, bar | Bidder to specify |  |  | Enter Responses Here!
Duty Cycle, % | The time a compressor can run without cooling off, longer duty cycle is preferred Bidder to specify |  |  | Enter Responses Here!
Noise Level, dB | Less than 70 dB |  |  | Enter Responses Here!
Noise Level, dB | Less than 60 dB |  |  | Enter Responses Here! | 3
Air Intake Filter | Required |  |  | Enter Responses Here!
Filter Size, µm | Bidder to specify |  |  | Enter Responses Here!
Drying Unit | Preferred, To prevent bacteria from growing |  |  | Enter Responses Here! | 3
Type | Desiccant or Membrane |  |  | Enter Responses Here! | 2
Pressure Dew Point, ˚C | Below -20˚C |  |  | Enter Responses Here! | 2
d. Vacuum Pumps
Type | Wet or Dry Vacuum Pump |  |  | Enter Responses Here!
Capacity | Bidder to ensure the number of vacuum pumps provided are sufficient to support all the required dental delivery units stipulated in the tender document |  |  | Enter Responses Here!
Maximum Number of Dental Delivery Units can be operated simultaneously | Bidder to specify |  |  | Enter Responses Here!
Maximum Suction Pressure, mbar (psi) | Bidder to specify |  |  | Enter Responses Here!
Speed, rpm | Bidder to specify |  |  | Enter Responses Here!
Noise Level, dB | Less than 70 dB |  |  | Enter Responses Here!
Noise Level, dB | Less than 60 dB |  |  | Enter Responses Here! | 3
General
Type of Disinfectant used for Infection Control | Free of corrosive ingredients |  |  | Enter Responses Here!
Compatible Agents | Bidder to specify |  |  | Enter Responses Here!
Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) High Speed Turbine Handpiece | 1 set per unit Bidder to specify part/order number for all handpiece provided |  |  | Enter Responses Here!
b) Low Speed Turbine Handpiece | 1 set per unit Bidder to specify part/order number for all handpiece provided |  |  | Enter Responses Here!
c) Handpiece Tips for Dental Ultrasonic Scalers | 1 set per unit Bidder to specify part/order number for all handpiece tips provided |  |  | Enter Responses Here! | 2
d) Handpiece Tips for Dental Air Abrasion | 1 set per unit Bidder to specify part/order number for all handpiece tips provided |  |  | Enter Responses Here! | 2
e) Foot Switch | 1 set per unit |  |  | Enter Responses Here!
Others, if any | Bidder to specify |  |  | Enter Responses Here!
START-UP CONSUMABLES, if any | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) Aluminium Oxide Powder for Dental Air Abrasion Unit (27µm) | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here! | 1
b) Aluminium Oxide Powder for Dental Air Abrasion Unit (50µm) | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
c) Air Compressor Unit Air Intake Filter | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
d) Air Compressors Drying Unit Filter | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. Dental Delivery Unit
Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here!
b. Dental Chair
Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here!
Weight Limit, kg | Bidder to specify |  |  | Enter Responses Here!
c. Medical Air Compressors
Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here!
d. Vacuum Pumps
Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here!
MECHANICAL & ELECTRICAL (M&E) REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. Dental Delivery Unit
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
b. Dental Chair
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
c. Medical Air Compressors
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
d. Vacuum Pumps
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
Others
Incoming Water Piping Opening Size, Inches | Bidder to specify |  |  | Enter Responses Here!
Water Pressure, Bar | Bidder to specify |  |  | Enter Responses Here!
Drainage Size, Inches | Bidder to specify |  |  | Enter Responses Here!
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | N/A |  |  | Enter Responses Here!
Compatibility with DICOM (Digital Imaging and Communications in Medicine) | N/A |  |  | Enter Responses Here!
Type of Networking/Connectivity Interface | N/A |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | N/A |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Bidder to specify if any | Bidder to provide installation, if applicable |  |  | Enter Responses Here!
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Equipment Room Readiness Requirement | To supply, install and complete commission of the room |  |  | Enter Responses Here!
Scope of Piping Works | Bidder to provide and integrate the Dental Delivery Units with Medical-Air Compressor and Central Vacuum Pump.
Incoming Water Piping Opening Size, Inches | Bidder to supply and install
Water Pressure, Bar | Bidder to supply and install
Drainage Size, Inches | Bidder to supply and install
Pre-Installation Requirements | All electrical power supply and installation to technical, control and equipment room
Air conditioning requirements | To supply, install and complete commissioning of required air conditioning in equipment room and technical room
Fire Alarm and Protection System | Bidder to supply and install
Temperature Requirements | The areas temperature ranges from 22-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed.
Humidity Requirements | The area's relative humidity ranges between 70% and 90%. Equipment must be operational in that range. If it is not please specify the range needed.
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a. Dental Delivery Unit
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 8
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!
b. Dental Chair
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 8
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!
c. Medical Air Compressors
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 3
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!
d. Vacuum Pumps
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 3
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Dental Chair (Complete System)
Include description and pricing for all accessories (e.g., <insert examples>, etc.) required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of Dental Chair (Complete System)
Include description and pricing for all consumable items (e.g.,<insert examples> etc.) required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Dental Chair (Complete System)
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!