## Infant Radiant Warmer
Infant Radiant Warmer |  |  |  |  | 42
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 17433 Warming Units, Patient, Radiant, Infant, Mobile Total Quantity: 1,884 | Equipment Location (Region): Quantity SOPHI a. West Indonesia: # units b. Central Indonesia: # units c. East Indonesia: # units | Description: Infant radiant warming units designed to provide thermal stability to an infant while allowing its transportation. These units typically consist of integral bassinet warming units mounted on casters, which may be equipped with brakes. Some warming units have an infant weighing scale, resuscitation equipment (suctioning and oxygen delivery), and vital signs measuring equipment (e.g., SpO2, pulse, airway pressure, body temperature). Mobile infant radiant warming devices are mostly used for internal transportation of infants from one area of the hospital to another and for warming and stabilizing a newborn infant in the delivery room after birth.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
As an radiant warmer for safely warming infants in an open environment | For Neonatal Intensive Care use |  |  | Enter Responses Here!
To perform resuscitation for newborn baby | For Obstetric Operating Rooms, Labor Delivery Rooms and Neonatal Intensive Care use |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Configuration | Mobile, Freestanding or with Integral Bassinet |  |  | Enter Responses Here!
Attached with | Oxygen tank, Regulators, Humidifier |  |  | Enter Responses Here!
Heating Element | Ceramic or Quartz Bidder to specify |  |  | Enter Responses Here! | 2
Heat Capacity, W | Preferred to be at least 500W Bidder to specify |  |  | Enter Responses Here! | 3
Mobile Stand
Height Adjustable | It is required the height of the basinet and heater element be adjustable |  |  | Enter Responses Here!
Height Adjustment - footpedal | It is preferred the height adjustment be controlled by footpedal. |  |  | Enter Responses Here!
Casters
Casters - number | The device shall have at least 4 casters |  |  | Enter Responses Here!
Casters Size, inches (cm) | At least 4'' Bidder to specify |  |  | Enter Responses Here!
Casters Lock | Required, at least 2 casters with lock to prevent unwanted movement |  |  | Enter Responses Here!
Tool Storage Drawers | Required. The device shall be available with drawers and trays. |  |  | Enter Responses Here!
Number of Drawers | Bidder to specify |  |  | Enter Responses Here!
Number of Trays/Shelves | Bidder to specify |  |  | Enter Responses Here!
IV Pole Holder | At least 1 Bidder to specify |  |  | Enter Responses Here!
IV Pole holder - more than 1 | It is preferred that the device accomodate more than one IV pole. |  |  | Enter Responses Here!
Physiologic Patient Monitor Holder | Required |  |  | Enter Responses Here!
Suction Canister Holder | Required |  |  | Enter Responses Here!
Gas Cylinder Holder | The device shall have at least two gas cylinder holders (one for air and one for oxygen). Bidder to specify |  |  | Enter Responses Here!
Bassinet
Integrated Bassinet | It is preferred to have an integrated bassinet. |  |  | Enter Responses Here!
Mattress Positioning Angle(Trenedelenberg) | The device should allow at least ten degrees of trendelelenberg and reverse trendelenberg. |  |  | Enter Responses Here!
Trendelenberg degrees. | Bidder to specify |  |  | Enter Responses Here!
Distance From Heat Source, cm | Bidder to specify |  |  | Enter Responses Here!
Resuscitation Unit
Resuscitation unit available | The device shall have a resuscitation unit. |  |  | Enter Responses Here!
Rususcitation mode | The device shall support the customer preferred resusciation mode, either bag or tee-piece. |  |  | Enter Responses Here! | 2
Compatibility | It is preferable that the resuscitator can support both resuscitation bags and Tee pieces. |  |  | Enter Responses Here! | 2
Other Resuscitation Modes. | Bidder to specify all resuscitation modes supported. |  |  | Enter Responses Here!
Gas Input Method - Wall Outlet | The resuscitator shall be able to be supplied from wall outlets |  |  | Enter Responses Here!
Gas Input Method - Gas Cylinder | The resuscitator shall be able to be supplied from gas tanks as a backup. |  |  | Enter Responses Here!
Gas Input Pressure Setting Range, psi | Bidder to specify |  |  | Enter Responses Here!
Oxygen-Air Blender
Oxygen - Air Blender | The resuscitation unit shall have a pressure controlled blended oxygen/air outlet |  |  | Enter Responses Here!
Oxygen Concentration - minimum | The blender should have a minimum setting of 21% and maximum setting of 100% |  |  | Enter Responses Here!
Oxygen Concentration Setting Range, % | Bidder to specify |  |  | Enter Responses Here!
Airway Pressure Range, cm H2O | Bidder to specify |  |  | Enter Responses Here!
Flow Range | The Flow range shall be adjustable from 0 to at least 15 LPM. |  |  | Enter Responses Here!
Flow Rate Setting Range, L/min | Bidder to specify |  |  | Enter Responses Here!
Suction/Vacuum Capability | The resuscitator shall provide suction capability from 0 to 150 mmHg. |  |  | Enter Responses Here!
Suction Range, mm Hg | Bidder to specify |  |  | Enter Responses Here!
Infant Scales
Integrated Scale | The device shall have an integrated scale |  |  | Enter Responses Here!
Weight Measurement Range, kg | Bidder to specify |  |  | Enter Responses Here!
Increment, g | Bidder to specify |  |  | Enter Responses Here!
Scale Functions
a) TARE | The scale shall have a tare mode. |  |  | Enter Responses Here!
b) HOLD | The scale shall have a hold mode. |  |  | Enter Responses Here!
c) RESET | The scale shall have a reset. |  |  | Enter Responses Here!
d) Auto-clear | Preferred |  |  | Enter Responses Here! | 1
e) Others | Bidder to specify other scale functions |  |  | Enter Responses Here!
X-Ray Tray | The warmer shall have an integrated, for emergency radiographic scanning |  |  | Enter Responses Here!
Cassette Tray Size, W x D (cm) | Bidder to specify |  |  | Enter Responses Here!
Examination Light | Required, for examining infants |  |  | Enter Responses Here!
Type of Light Source | LED, White-light |  |  | Enter Responses Here!
Illumination, Lux | Bidder to specify |  |  | Enter Responses Here!
Operation Modes for Temperature Control | Required; Both Servo-controlled (Automatic) mode and Manual mode is available |  |  | Enter Responses Here!
a) Servo-controlled (Automatic) Mode | Heating element should turn on and off in response to changes in the infant's skin temperature. |  |  | Enter Responses Here!
Temperature Range, °C | Within 34°C to 38°C. Should not exceed 39°C. Bidder to specify |  |  | Enter Responses Here!
Temperature Increment, °C | Bidder to specify |  |  | Enter Responses Here!
b) Manual (non-servo) Mode
Manual mode - available | The warmer shall have an adjustable from 0% to 100% in 5% increment |  |  | Enter Responses Here!
Manual Timeout | The warmer shall have a manual mode timeout that turns heater power off within 30 minutes if not reset. This prevents inadvertant overheating while being left unattended. |  |  | Enter Responses Here!
Manual mode timeout | Bidder to specify the time that the uint can remain in manual mode before the heater is turned off. |  |  | Enter Responses Here!
Manual mode time alarm | The warmer shall alarm when left in manual mode. This should occur prior to heater cutout |  |  | Enter Responses Here!
Manual mode time alarm | Bidder to specify the time that the unit will alarm when left in manual mode. |  |  | Enter Responses Here!
Control Panel | Touch panel |  |  | Enter Responses Here!
Display Type | LCD to LED |  |  | Enter Responses Here!
Displayed Information | a) Set-point temperature |  |  | Enter Responses Here!
| b) Skin temperature |  |  | Enter Responses Here!
| c) Warming time / Timer |  |  | Enter Responses Here!
| d) Heater output level |  |  | Enter Responses Here!
| e) Power on/off status |  |  | Enter Responses Here!
| f) Others, Bidder to specify |  |  | Enter Responses Here!
Timer | Required, to limit the heating time, and an alarm to prompt reassessment of the infant's status |  |  | Enter Responses Here!
Apgar Timer | Required |  |  | Enter Responses Here!
Alarms | Both audible and visual indicators |  |  | Enter Responses Here!
Alarms Conditions | a) High and low skin temperatures |  |  | Enter Responses Here!
| b) Sensor disconnection or failure |  |  | Enter Responses Here!
| c) Fan failure |  |  | Enter Responses Here!
| d) Power failure |  |  | Enter Responses Here!
| e) Others, bidder to specify |  |  | Enter Responses Here!
Trigger Alarms for Skin Temperature from Setpoint, ±°C | Within 0.5˚ to 1˚C from set-point |  |  | Enter Responses Here!
Tones of Alarm Volume | Variable or fixed |  |  | Enter Responses Here!
Self-Check Feature | Required, to verify correct function of the electronics and components and to ensure proper functioning |  |  | Enter Responses Here!
Backup Thermostat | Required, a system to deactivate the heater when the primary controller malfunction |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Skin Temperature Probe | 1 set per unit |  |  | Enter Responses Here!
IV pole | 1 pc per unit |  |  | Enter Responses Here!
Suction Canisters, 500mL | 1 pc per unit |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
START-UP CONSUMABLES, if any | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Eye Patches or Eye Masks (For Infant) | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
Mattress Cover | Bidder to supply quantity for 100 typical patient procedures. |  |  | Enter Responses Here!
Other Consumables | Bidder to specify if there are other consumables needed |  |  | Enter Responses Here!
Consumables needed | Bidder to specify itemized quantity for 100 typical patient procedures |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Unit Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Bassinet Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Bassinet Height Range, H, cm | Bidder to specify |  |  | Enter Responses Here!
Mattress Thickness, H (cm) | Bidder to specify |  |  | Enter Responses Here!
Unit Weight, kg | Bidder to specify |  |  | Enter Responses Here!
ELECTRICAL REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Electrical Requirement, V | 230 ± 10 V AC, 50Hz |  |  | Enter Responses Here!
Power Plug Type | Type F power plug, in complicance with Indonesia requirements. |  |  | Enter Responses Here!
Power Consumption, kW | Bidder to specify highest power consumption |  |  | Enter Responses Here!
UPS Requirement | Preferred; Bidder to specify availability |  |  | Enter Responses Here! | 3
Estimated Operation Time, minutes | At least 30 minutes |  |  | Enter Responses Here! | 2
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Preferred |  |  | Enter Responses Here! | 3
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Interface with Information System | Preferred, e.g: HIS, RIS, EMR, etc |  |  | Enter Responses Here! | 3
Type of Networking/Connectivity Interface | Wireless or LAN-Port, or both Bidder to specify |  |  | Enter Responses Here!
Compatible Data Management System (Middleware) | Bidder to specify |  |  | Enter Responses Here!
Data Port | Data port must be activated (e.g: RS232, USB) to allow download of data logs from device and service |  |  | Enter Responses Here! | 1
Communication Protocol | Bidder to specify |  |  | Enter Responses Here!
System Upgradeable in Future | Software and hardware updates and upgrades at no cost |  |  | Enter Responses Here!
Remote Support Capability | Preferred; Off site diagnosis of technical issues and remote fixes. |  |  | Enter Responses Here! | 3
Cybersecurity Considerations
Installation of Current Release Software Version | Bidder to provide |  |  | Enter Responses Here!
Anti-malware or Anti-virus, included or compatible | Preferred. |  |  | Enter Responses Here! | 2
Deactivation of unnecessary open communication channels (e.g.: USB ports) | Preferred; Ability to disable (software and hardware) |  |  | Enter Responses Here! | 1
Confidentiality and Security Measures | Bidder to specify if applicable |  |  | Enter Responses Here!
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A |  |  | Enter Responses Here!
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A |  |  | Enter Responses Here!
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Oxygen | Bidder to provide installation. |  |  | Enter Responses Here!
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Temperature Requirements | The areas temperature ranges from 22-32 C. Equipment must be operational in that temperature range. If it is not please specify the range needed. |  |  | Enter Responses Here!
Humidity Requirements | The area's relative humidity ranges between 70% and 90%.Equipment must be operational in that range. If it is not please specify the range needed. |  |  | Enter Responses Here!
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 6 and 7, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 14
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 10 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Infant Radiant Warmer
Include description and pricing for all accessories required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of Infant Radiant Warmer
Include description and pricing for all consumable items required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Infant Radiant Warmer
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!