## Nebulizers
Nebulizers |  |  |  |  | 38 | <--- Total score
Bidder shall complete this form to provide responses about their proposed medical equipment. Each response by Bidder shall be supported with reference document,e.g: Catalogue, Equipment Technical Data Sheet, etc., and each reference shall be highlighted and/or labelled.
UMDNS: 12712 Nebulizers Total Quantity: 1,917 | Equipment Location (Region): Quantity SOPHI a. Indonesia (Nationwide): 2,939 units | Description: Devices designed to provide aerosolized medication and/or humidification to the patient airways. These devices produce both vapor and finely dispersed water droplets in the liquid phase (aerosol mist); they are typically used with a mask or tent or in a breathing circuit. Most nebulizers are either pneumatic (driven by compressed air) or ultrasonic. Most nebulizers include a heating element to provide a warm, humid aerosol.
PROPOSED DEVICE INFORMATION | REQUIREMENTS | BIDDER'S RESPONSE
Brand/Manufacturer Name | Bidder to specify | Enter Responses Here!
Model Name | Bidder to specify | Enter Responses Here!
Country of Origin | Bidder to specify | Enter Responses Here!
Year of 1st Production for Model Proposed | Bidder to specify | Enter Responses Here!
Brochure/Catalogue | Bidder shall provide product brochure in English. Both Soft copy and Hard copy are required. | Enter Responses Here!
Nomor Izin Edar (NIE) - Indonesia Product Approval License | NIE will not be a prerequisite at bidding stage, yet required for contract award. Bidder to specify whether NIE has already been alloted, as detailed in the bidding document | Enter Responses Here!
U.S. FDA Clearance, CE Mark (Conformité Européenne mark) (MDD), or other health, safety and environmental standard adherence | Mandatory if NIE not allotted as of yet. Bidder to provide evidence | Enter Responses Here!
Picture of Model Proposed | Bidder to attach a picture of model proposed | Enter Responses Here!
CLINICAL APPLICATIONS/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Provide aerosolized medication and/or humidification for treating severe diseases of the airway | For adults, pediatrics and neonates |  |  | Enter Responses Here!
TECHNICAL SPECIFICATIONS/FEATURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Type | Portable |  |  | Enter Responses Here! | 5
No. of Nozzles | Bidder to specify |  |  | Enter Responses Here!
Total output rate, mg/min | > 200 mg/min |  |  | Enter Responses Here! | 5
Median Mass Aerodynamic Diameter (MMAD) of particles | < 2.5 micron |  |  | Enter Responses Here! | 5
Medication capacity | ≥ 6 ml |  |  | Enter Responses Here! | 5
Noise level | ≤ 55 dB |  |  | Enter Responses Here! | 3
Complete set | Includes medication reservoir, nozzle, tubing, inspiratory valve, and mask. |  |  | Enter Responses Here!
Type of Disinfectant used for Infection Control | Free of corrosive ingredients |  |  | Enter Responses Here!
Compatible Agents | Bidder to specify |  |  | Enter Responses Here!
Cleaning Instruction/Guideline for Clinicians/Users | Bidder to provide |  |  | Enter Responses Here!
ACCESSORIES REQUIRED FOR EACH CLINICAL APPLICATION/PROCEDURES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
a) Adult mask | 1 set per unit |  |  | Enter Responses Here!
b) Pediatric mask | 1 set per unit |  |  | Enter Responses Here!
c) Neonate mask | 1 set per unit |  |  | Enter Responses Here!
Others | Bidder to specify |  |  | Enter Responses Here!
DEVICE PHYSICAL ATTRIBUTES | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Dimension, H x W x D, cm | Bidder to specify |  |  | Enter Responses Here!
Weight, kg | Bidder to specify |  |  | Enter Responses Here!
ELECTRICAL REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Electrical Requirement, V | 230 ± 10 V AC, 50Hz, Single-Phase |  |  | Enter Responses Here!
Power Plug Type | Hospital Grade Type F power plug, in compliance with Indonesia requirements |  |  | Enter Responses Here!
Power Consumption, W | Bidder to specify highest power consumption |  |  | Enter Responses Here!
Battery Powered | Rechargeable |  |  | Enter Responses Here!
Battery Operating Time, hr | Bidder to specify |  |  | Enter Responses Here!
Battery Recharging Time, hr | Bidder to specify |  |  | Enter Responses Here!
UPS Requirement | Bidder to specify |  |  | Enter Responses Here!
Estimated Operation Time, minutes | At least 30 minutes |  |  | Enter Responses Here!
Compliant with IEC 60601 - Safety of Medical Electrical Equipment | Required. Bidder to provide certificate. |  |  | Enter Responses Here!
Compliant with IEC Standards for Energy Efficiency | Bidder to specify relevant IEC Standards |  |  | Enter Responses Here!
Certifications related to Energy Efficiency | Bidder to specify |  |  | Enter Responses Here! | 3
IT REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ARTIFICIAL INTELLIGENCE (AI) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
TELEMEDICINE (TM) CAPABILITITES, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
MEDICAL GASES REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
ROOM REQUIREMENTS, where applicable | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
N/A | N/A
MAINTENANCE & SERVICE REQUIREMENTS | REQUIREMENTS | BIDDER'S RESPONSE
|  | Yes (√) | No (x) | Bidder to enter responses about their device
Acceptance Testing & Commissioning | Yes, inclusive of vendor representative attendance, with test equipments and supplies Bidder must perform Acceptance Testing & Commissioning that includes Electrical Safety Testing using calibrated test equipment and tools, and procedural checklist. Bidder shall inspect and test all medical equipment supplied to hospital in accordance to manufacturer's recommended procedures. |  |  | Enter Responses Here!
User Training | Required. Twice, within 4 weeks upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Duration of User Training Session (Time) | Bidder to specify the duration of training session |  |  | Enter Responses Here!
Operator Manual (Instruction For Use) and Service Manual | Required. Bidder to provide in Hardcopy and/or Softcopy upon successful Acceptance Testing & Commissioning. |  |  | Enter Responses Here!
Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Warranty Duration: Minimum 2 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Extention of Warranty Duration | To commence upon successful Acceptance Testing & Commissioning. Extension Warranty Duration: Year 3 and 4, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here! | 12
Manufacturer Recommended Annual Preventive Maintenance (Kalibrasi) Interval | Bidder to specify, by number of times a year, minimum 2 times per year |  |  | Enter Responses Here!
Comprehensive Service Contract Duration, Post Warranty | To commence upon end of Warranty Duration. Comprehensive Service Contract Duration: Minimum 5 years, inclusive of annual preventive maintenance and repair services; inclusive of all labor, parts, logistics and transportation |  |  | Enter Responses Here!
Response Time, Remote Technical Support (Phone/Email) | Within 1 hour upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Response Time, On-site Technical Support | Within 5 calendar days upon notification from user, inclusive of weekends and public holidays Applicable for both Warranty Duration and Comprehensive Service Contract Duration. |  |  | Enter Responses Here!
Turnaround time for corrective maintenance | less than 5 calendar days (with penalty system) |  |  | Enter Responses Here!
Uptime Guarantee, Annually (days) | ≥95% |  |  | Enter Responses Here!
Assured Availability of Spare Parts | Minimum 5 years post Acceptance Testing and Commissioning Date |  |  | Enter Responses Here!
Installation Base (Specific to this model) | Please provide installation base (hospital name and location), year of installation |  |  | Enter Responses Here!

## Hazards and Recalls
Notification of Equipment Hazards and/or Recalls
Please list down the Hazards and/or Recalls of the equipment that you proposed and the field corrective actions, if any.
List (Reported Hazards, Alerts, RECALL): | Please provide clarification of action taken:
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!
Enter Responses Here! | Enter Responses Here!

## Proposed Accessories and Price
Proposed Accessories and Price (5 years)
Procurement of Nebulizers
Include description and pricing for all accessories (e.g., <insert examples>, etc.) required for a typical procedure and offered by your company.
Brand | Model | Accessories Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Consumables and Price
Proposed Consumables and Price (5 years)
Procurement of Nebulizers
Include description and pricing for all consumable items (e.g.,<insert examples> etc.) required for a typical procedure and offered by your company. Provide total consumable pricing for 100 typical patient procedures.
Brand | Model | Consumables Name/Description | Catalogue No. | UoM/Quantity | Unit Price ($) | Total Price ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!
0
0
0
0
0
0
0
0
0
0
0
0
0
0
|  | TOTAL |  | 0 | 0 | 0
|  | Discount ($), if any |  |  |  | 0
|  | GRAND TOTAL PRICE ($) |  |  |  | 0

## Proposed Post Warranty Service
Proposed Post - Warranty, Maintenance Service Contract (Year 6 to 10)
Procurement of Nebulizers
Brand | Model | UoM/Quantity | PPM Frequency (times/year) | Comprehensive Service Contract
Shall include: - Working hours - Labor charge and cost incurred in transportation & accomodation - Spare parts
|  |  |  | Details of Service Contract | Total for Year 6 ($) | Total for Year 7 ($) | Total for Year 8 ($) | Total for Year 9 ($) | Total for Year 10 ($)
Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here! | Enter Response Here!